## Document information
Revision | Date | By | Changes
1 | 2014-08-08 | Oksana Ivashkevych | Original version
| 2015-03-01 | Oksana Ivashkevych | Current version

## LBD
Total motor Axis count | GigE CCD cams | Detectors x2 Pilatus300M, x1 Pilatus1M | Pt100 | K type TC | Digital Inputs | Digital Outputs | Analog IO Inputs | Analog IO Outputs
108 | 5 | 3 | 21 | 16 | 99 | 25 | 2 | 1
Some important dates for contracted out components' arrival
WBM System | 2015-06-01
Mono System | 2015-06-01
KB Mirror System | 2015-06-01
Flight Path End Station | 2015-07-01

## Component summary
Beamline Identifier | XF05ID |  | Driver Tested | Database tested | Application/GUIs
Component | Hutch | Location
LiX FOE Optics Hutch A | FE
Beam-viewing screen #1 | A |  | V | V
Fixed Mask | A
Bremsstrahlung collimator 1 | A
White Beam Mirror | A
Beam-viewing screen #2 | A |  | V | V
White Beam Stop | A
Bremsstrahlung collimator 2 | A
Dual Multilayer Monochromator | A
Si(111) Monochromator (DCM/CCM) | A
Pink Beam Stop | A
Bremsstrahlung collimator 3 | A
Beam-viewing screen #3 | A |  | V | V
Slits S0 | A
KB mirror system | A
Beam-viewing screen #4 | A |  | V | V
Photon Shutter | A
A
Secondry Source SOE Hutch B
Beam transport | B
Scintilator #5
Beam Position Monitor (BPM) | B
Secondary Source Chamber | B
Secondary Source Aperture(SSA) | B |  | V | V
Attenuator | B |  | V
Visual Beam Monitor (VBM) | B
Shutter | B
Alternative SSA | B
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | C
Secondary focusing working distance | C
Harmonic Regection Mirror HRM1 | C |  | V | V
Transfocator CRLs | C
Harmonic Regection Mirror HRM2 | C |  | V | V
Divergence Defining Aperture (DDA) | C
Beam Position Monitor (BPM) | C
Guard Slits 1 in Air | C
Experimental module (user defined configuration) | C
In-air transmission measurement module
Coarse-resolution scanning | C
Fine-resolution scanning | C
Microscope | C
Guard Slits 2 in Air | C
In-vaccum GISAXS/GID module
Hexapod | C
Guard Slits 2 in VAC | C |  | V | V
Sample handling
Solution sample handler | C
Staubli TX40 robot | C
Detector system
WAXS chamber | C
WAXS detector 1 | C
WAXS detector 2 | C
WAXS beamstop | C |  | V
SAXS detector | C
SAXS beam stop | C |  | V
Flight Path module | C
Note on A55: This will include multiple stages; hope to have integrated controller
Note on A56: This has its own controller

## Component naming
Beamline Identifier | 16ID
Component | Hutch | Primary subsystem | Primary instance | Primary ID | Secondary subsystem | Secondary instance | Secondary ID | Component name | Component instance | Component ID | Full component name
LiX FOE Optics Hutch A
Beam-viewing screen #1 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Scr | SCN1 | Scr:SCN1 | XF:16IDA-OP:LiX{Scr:SCN1}
Fixed Mask | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Msk |  | Msk | XF:16IDA-OP:LiX{Msk}
Bremsstrahlung collimator 1 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | BS | 1 | BS:1 | XF:16IDA-OP:LiX{BS:1}
White Beam Mirror | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Mir | WBM | Mir:WBM | XF:16IDA-OP:LiX{Mir:WBM}
Beam-viewing screen #2 | A | XF | 16IDA | XF:16IDA | BI:LiX |  | BI:LiX | Mon | SCN2 | Mon:SCN2 | XF:16IDA-BI:LiX{Mon:SCN2}
White Beam Stop | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | BS | White | BS:White | XF:16IDA-OP:LiX{BS:White}
Bremsstrahlung collimator 2 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | BS | 2 | BS:2 | XF:16IDA-OP:LiX{BS:2}
Dual Multilayer Monochromator | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Mono | ML | Mono:ML | XF:16IDA-OP:LiX{Mono:ML}
Si(111) Monochromator (DCM/CCM) | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Mono | Si | Mono:Si | XF:16IDA-OP:LiX{Mono:Si}
Pink Beam Stop | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | BS | Pink | BS:Pink | XF:16IDA-OP:LiX{BS:Pink}
Bremsstrahlung collimator 3 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | BS | 3 | BS:3 | XF:16IDA-OP:LiX{BS:3}
Beam-viewing screen #3 | A | XF | 16IDA | XF:16IDA | BI:LiX |  | BI:LiX | Scr | SCN3 | Scr:SCN3 | XF:16IDA-BI:LiX{Scr:SCN3}
Slits S0 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Slt | S0 | Slt:S0 | XF:16IDA-OP:LiX{Slt:S0}
KB mirror system | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Mir | KB | Mir:KB | XF:16IDA-OP:LiX{Mir:KB}
Beam-viewing screen #4 | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Mon | SCN4 | Mon:SCN4 | XF:16IDA-OP:LiX{Mon:SCN4}
Photon Shutter | A | XF | 16IDA | XF:16IDA | OP:LiX |  | OP:LiX | Sh |  | Sh | XF:16IDA-OP:LiX{Sh}
Secondry Source SOE Hutch B
Beam transport
Scintilator #5 | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Mon | SCN5 | Mon:SCN5 | XF:16IDB-SS:LiX{Mon:SCN5}
Beam Position Monitor (BPM) | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Mon | BPM1 | Mon:BPM1 | XF:16IDB-SS:LiX{Mon:BPM1}
Secondary Source Chamber | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Chm | SS | Chm:SS | XF:16IDB-SS:LiX{Chm:SS}
Secondary Source Aperture(SSA) | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Slt | SSA | Slt:SSA | XF:16IDB-SS:LiX{Slt:SSA}
Attenuator | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Attn |  | Attn | XF:16IDB-SS:LiX{Attn}
Visual Beam Monitor (VBM) | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Mon | VBM | Mon:VBM | XF:16IDB-SS:LiX{Mon:VBM}
Shutter | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Sh |  | Sh | XF:16IDB-SS:LiX{Sh}
Alternative SSA | B | XF | 16IDB | XF:16IDB | SS:LiX |  | SS:LiX | Slt | aSSA | Slt:aSSA | XF:16IDB-SS:LiX{Slt:aSSA}
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Chm | SF | Chm:SF | XF:16IDC-SF:LiX{Chm:SF}
Secondary focusing working distance | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Stg | WD | Stg:WD | XF:16IDC-SF:LiX{Stg:WD}
Harmonic Regection Mirror HRM1 | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Mir | HRM1 | Mir:HRM1 | XF:16IDC-SF:LiX{Mir:HRM1}
Transfocator CRLs | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Trf |  | Trf | XF:16IDC-SF:LiX{Trf}
Harmonic Regection Mirror HRM2 | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Mir | HRM2 | Mir:HRM2 | XF:16IDC-SF:LiX{Mir:HRM2}
Divergence Defining Aperture (DDA) | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Slt | DDA | Slt:DDA | XF:16IDC-SF:LiX{Slt:DDA}
Beam Position Monitor (BPM) | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Mon | BPM2 | Mon:BPM2 | XF:16IDC-SF:LiX{Mon:BPM2}
Guard Slits 1 in Air | C | XF | 16IDC | XF:16IDC | SF:LiX |  | SF:LiX | Slt | G1 | Slt:G1 | XF:16IDC-SF:LiX{Slt:G1}
Experimental module (user defined configuration) | C | XF | 16IDC | XF:16IDC | SF:LiX |  | ES |  | ES | :ES | XF:16IDC-ES{:ES}
In-air transmission measurement module
Coarse-resolution scanning | C | XF | 16IDC | XF:16IDC | ES:LiX | InAir | ES:LiX:InAir | Stg | CR | Stg:CR | XF:16IDC-ES:LiX:InAir{Stg:CR}
Fine-resolution scanning | C | XF | 16IDC | XF:16IDC | ES:LiX | InAir | ES:LiX:InAir | Stg | FR | Stg:FR | XF:16IDC-ES:LiX:InAir{Stg:FR}
Microscope | C | XF | 16IDC | XF:16IDC | ES:LiX | InAir | ES:LiX:InAir | Mscp |  | Mscp | XF:16IDC-ES:LiX:InAir{Mscp}
Guard Slits 2 in Air | C | XF | 16IDC | XF:16IDC | ES:LiX | InAir | ES:LiX:InAir | Slt | G2 | Slt:G2 | XF:16IDC-ES:LiX:InAir{Slt:G2}
In-vaccum GISAXS/GID module
Hexapod | C | XF | 16IDC | XF:16IDC | ES:LiX | GI | ES:LiX:GI | Stg | Hex | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex}
Guard Slits 2 in VAC | C | XF | 16IDC | XF:16IDC | ES:LiX | GI | ES:LiX:GI | Slt | G2 | Slt:G2 | XF:16IDC-ES:LiX:GI{Slt:G2}
Sample handling
Solution sample handler | C | XF | 16IDC | XF:16IDC | ES:LiX |  | ES:LiX | Rbt | Sol | Rbt:Sol | XF:16IDC-ES:LiX{Rbt:Sol}
Staubli TX40 robot | C | XF | 16IDC | XF:16IDC | ES:LiX |  | ES:LiX | Rbt | Staubli | Rbt:Staubli | XF:16IDC-ES:LiX{Rbt:Staubli}
Detector system
WAXS chamber | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | Chm | WAXS | Chm:WAXS | XF:16IDC-DT:LiX{Chm:WAXS}
WAXS detector 1 | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | Det | WAXS1 | Det:WAXS1 | XF:16IDC-DT:LiX{Det:WAXS1}
WAXS detector 2 | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | Det | WAXS2 | Det:WAXS2 | XF:16IDC-DT:LiX{Det:WAXS2}
WAXS beamstop | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | BS | WAXS | BS:WAXS | XF:16IDC-DT:LiX{BS:WAXS}
SAXS detector | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | Det | SAXS | Det:SAXS | XF:16IDC-DT:LiX{Det:SAXS}
SAXS beam stop | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX | BS | SAXS | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS}
Flight Path module | C | XF | 16IDC | XF:16IDC | DT:LiX |  | DT:LiX |  | ??? | :??? | XF:16IDC-DT:LiX{:???}

## PLC IO summary
Component | Component ID | Hutch | Location | Temperature sensors |  | Digital IO |  |  |  | Analog IO
|  |  | (mm from source) | Pt100 | K type TC | Inputs | Type | Outputs | Type | Inputs | Type | Outputs | Type
Totals |  |  |  | 21 | 16 | 99 |  | 25 |  | 2 |  | 1
LiX FOE Optics Hutch A
Beam-viewing screen #1 | Scr:SCN1 | A |  | 2 |  | 2 | Dry contact | 1 | 24V solenoid
Fixed Mask | Msk | A |  | 2
Bremsstrahlung collimator 1 | BS:1 | A
White Beam Mirror | Mir:WBM | A |  | 4 | 3
Beam-viewing screen #2 | Mon:SCN2 | A |  | 2 |  | 2 | Dry contact | 1 | 24V solenoid
White Beam Stop | BS:White | A |  | 0 | 4
Bremsstrahlung collimator 2 | BS:2 | A
Dual Multilayer Monochromator | Mono:ML | A |  | 2
Si(111) Monochromator (DCM/CCM) | Mono:Si | A |  | 9 | 4 | 2 | Dry contact |  |  |  |  | 1 |  | for slow feedback system
Pink Beam Stop | BS:Pink | A |  | 0 | 2
Bremsstrahlung collimator 3 | BS:3 | A |  | 0 | 0
Beam-viewing screen #3 | Scr:SCN3 | A |  | 0 | 0
Slits S0 | Slt:S0 | A |  | 0 | 0
KB mirror system | Mir:KB | A |  | 0 | 3
Beam-viewing screen #4 | Mon:SCN4 | A |  | 0 | 0
Photon Shutter | Sh | A
Subtotal |  |  |  | 21 | 16 | 6 |  | 2 |  | 0 |  | 1
Secondry Source SOE Hutch B
Beam transport
Scintilator #5 | Mon:SCN5 | B
Beam Position Monitor (BPM) | Mon:BPM1 | B
Secondary Source Chamber | Chm:SS | B |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | rouphing pump
Secondary Source Aperture(SSA) | Slt:SSA | B
Attenuator | Attn | B
Visual Beam Monitor (VBM) | Mon:VBM | B
Shutter | Sh | B
Alternative SSA | Slt:aSSA | B
General purpose IO |  | B
Subtotal |  |  |  | 0 | 0 | 2 |  | 1 |  | 0 |  | 0
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | Chm:SF | C |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | rouphing pump
Secondary focusing working distance | Stg:WD | C
Harmonic Regection Mirror HRM1 | Mir:HRM1 | C
Transfocator CRLs | Trf | C |  |  |  | 27 |  | 9
Harmonic Regection Mirror HRM2 | Mir:HRM2 | C | [7 empty] | 2 |  |  |  | from LVDT sensor, additional module integration
Divergence Defining Aperture (DDA) | Slt:DDA | C
Beam Position Monitor (BPM) | Mon:BPM2 | C
Guard Slits 1 in Air | Slt:G1 | C
Experimental module (user defined configuration) | :ES | C |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | rouphing pump
In-air transmission measurement module
Coarse-resolution scanning | Stg:CR | C
Fine-resolution scanning | Stg:FR | C
Microscope | Mscp | C
Guard Slits 2 in Air | Slt:G2 | C
In-vaccum GISAXS/GID module
Hexapod | Stg:Hex | C
Guard Slits 2 in VAC | Slt:G2 | C
Sample handling
Solution sample handler | Rbt:Sol | C
Staubli TX40 robot | Rbt:Staubli | C
Detector system
WAXS chamber | Chm:WAXS | C |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | rouphing pump
WAXS detector 1 | Det:WAXS1 | C
WAXS detector 2 | Det:WAXS2 | C
WAXS beamstop | BS:WAXS | C
SAXS detector | Det:SAXS | C
SAXS beam stop | BS:SAXS | C
General purpose IO |  | C
Subtotal |  |  |  | 0 | 0 | 33 |  | 12 |  | 2 |  | 0
Vacuum system
XF:16IDA-VA{FE-GV:1} |  | A | 0 |  |  | 0 |  | 0 | [5 empty] | Handled by front end
XF:16IDA-VA{Scr:SCN1-IP:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{Scr:SCN1-CCG:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{Scr:SCN1-TCG:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{BS:1-IP:1} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{BS:1-CCG:1} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{BS:1-TCG:1} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{BS:1-GV:1} |  | A | 28.2 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mir:WBM-IP:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-CCG:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-TCG:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-GV:1} |  | A | 29.1 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mono:ML-IP:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:ML-CCG:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:ML-TCG:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:ML-IP:1} |  | A | 31.5 |  |  | 1
XF:16IDA-VA{Mono:ML-CCG:1} |  | A | 31.5 |  |  | 1
XF:16IDA-VA{Mono:Si-TCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:Si-GV:1} |  | A | 32.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mono:Si-IP:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:Si-CCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:Si-TCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:Si-GV:2} |  | A | 32.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Scr:SCN3-IP:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{Scr:SCN3-CCG:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{Scr:SCN3-TCG:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{Scr:SCN3-GV:1} |  | A | 36.3 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mir:KB-IP:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-CCG:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-TCG:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-GV:1} |  | A | 37.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mon:SCN4-IP:1} |  | A | 39 |  |  | 1
XF:16IDA-VA{Mon:SCN4-CCG:1} |  | A | 39 |  |  | 1
XF:16IDA-VA{Mon:SCN4-TCG:1} |  | A | 39 |  |  | 1
XF:16IDB-VA{Mon:SCN5-IP:1} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{Mon:SCN5-CCG:1} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{Mon:SCN5-TCG:1} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{Mon:SCN5-GV:1} |  | B | 50.1 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDB-VA{Chm:SS-TCG:1} |  | B | 51 |  |  | 1
XF:16IDB-VA{Chm:SS-TCG:2} |  | B | 51 |  |  | 1
XF:16IDC-VA{Chm:SF-GV:1} |  | C | 51 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{Chm:SF-TCG:1} |  | C | 58.875 |  |  | 1
XF:16IDC-VA{Chm:SF-GV:2} |  | C | 58.9 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{:ES-TCG:1} |  | C | 59.5 |  |  | 1
XF:16IDC-VA{:ES-GV:1} |  | C | 59.5 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{Chm:WAXS-TCG:1} |  | C | 60 |  |  | 1
Equipment identification signal |  | C | 60 |  |  | 3
Subtotal |  |  |  | 0 | 0 | 58 |  | 10 |  | 0 |  | 0
Cooling water
Beam-viewing screen #1 | Scr:SCN1 | A | Roof | [6 empty] | 1 | 4-20 mA
Fixed Mask | Msk | A | Roof | [6 empty] | 1 | 4-20 mA
White Beam Stop | BS:White | A | Roof | [6 empty] | 1 | 4-20 mA
Beam-viewing screen #2 | Mon:SCN2 | A | Roof | [6 empty] | 1 | 4-20 mA
Pink Beam Stop | BS:Pink | A | Roof | [6 empty] | 1 | 4-20 mA
Subtotal |  |  |  | 0 | 0 | 0 |  | 0 |  | 5 |  | 0
Total |  |  |  | 21 | 16 | 99 |  | 25 |  | 2 |  | 1
Other (So far irrelevant) Info below is from SRX
Synchrotron beam not inhibited by PPS | [5 empty] | 1 | Scott Buda to define
Synchrotron beam not inhibited by EPS | [7 empty] | 1 | Scott Buda to define
Cooling water supply isolation valve |  | A | Upstream |  |  | 2 |  | 2
Cooling water return isolation valve |  | A | Upstream |  |  | 2 |  | 2
Leak detection system |  | A | ? |  |  | 2?
Cooling water supply temperature |  | A | Upstream | 1
Cooling water return temperature |  | A | Upstream | 1
Cryocooler |  | A | Mezzanine |  |  | 1
Cryocooler disable |  | A | Mezzanine |  |  |  |  | 1
A Hutch O2 sensor |  | A | ? |  |  | 1
D Hutch O2 sensor |  | D | ? |  |  | 1
N2 system disable |  | A | ? |  |  |  |  | 1
Motion controller disable |  | A | RG:A2 |  |  |  |  | 4
Motion controller disable |  | D | RG:D1? |  |  |  |  | 4
D Hutch endstation motion interlocks |  | D | 65000 |  |  | 20 | Dry contact | 10 | 24 VDC solenoid
A Hutch PPS |  | A |  |  |  | 3 |  | 3
B Hutch PPS |  | B |  |  |  | 3 |  | 2
D Hutch PPS |  | D |  |  |  | 3 |  | 2
Subtotal |  |  |  | 2 | 0 | 37 |  | 32 |  | 0 |  | 0
Total |  |  |  | #REF! | #REF! | #REF! |  | #REF! |  | #REF! |  | #REF!
Note on A123: Ivashkevych, Oksana: this is to identify which component is installed, up to 2^3; some logig will be involved, you have to first disconnect A before connecting B

## Vacuum components
Beamline Identifier | 16ID
Component | Hutch | Location (meters) | Primary subsystem | Primary instance | Primary ID | Secondary subsystem | Branch | Secondary ID | Primary device name | Device name | Device instance | Device ID | Full component name | Associated component | Controller | Channel | Cable number
LiX Hutch A
Gate valve | A |  | XF | 16IDA | XF:16IDA | VA |  | VA | FE | GV | 1 | GV:1 | XF:16IDA-VA{FE-GV:1}
Ion pump | A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN1 | IP | 1 | IP:1 | XF:16IDA-VA{Scr:SCN1-IP:1} |  | XF:16IDA-CT{IPC:1} | 1 | 74
Cold cathode gauge | A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN1 | CCG | 1 | CCG:1 | XF:16IDA-VA{Scr:SCN1-CCG:1} |  | XF:16IDA-CT{VGC:1} | 1 | 47
Pirani gauge | A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN1 | TCG | 1 | TCG:1 | XF:16IDA-VA{Scr:SCN1-TCG:1} |  | XF:16IDA-CT{VGC:1} | 5 | 99
Ion pump | A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | BS:1 | IP | 1 | IP:1 | XF:16IDA-VA{BS:1-IP:1} |  | XF:16IDA-CT{IPC:1} | 2 | 69
Cold cathode gauge | A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | BS:1 | CCG | 1 | CCG:1 | XF:16IDA-VA{BS:1-CCG:1} |  | XF:16IDA-CT{VGC:1} | 3 | 42
Pirani gauge | A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | BS:1 | TCG | 1 | TCG:1 | XF:16IDA-VA{BS:1-TCG:1} |  | XF:16IDA-CT{VGC:1} | 6 | 100
Gate valve | A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | BS:1 | GV | 1 | GV:1 | XF:16IDA-VA{BS:1-GV:1}
Ion pump | A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | IP | 1 | IP:1 | XF:16IDA-VA{Mir:WBM-IP:1} |  | XF:16IDA-CT{IPC:2} | 1 | 77
Cold cathode gauge | A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | CCG | 1 | CCG:1 | XF:16IDA-VA{Mir:WBM-CCG:1} |  | XF:16IDA-CT{VGC:2} | 1 | 50
Pirani gauge | A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | TCG | 1 | TCG:1 | XF:16IDA-VA{Mir:WBM-TCG:1} |  | XF:16IDA-CT{VGC:2} | 5 | 107
Gate valve | A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | GV | 1 | GV:1 | XF:16IDA-VA{Mir:WBM-GV:1}
Ion pump | A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:ML | IP | 1 | IP:1 | XF:16IDA-VA{Mono:ML-IP:1} |  | XF:16IDA-CT{IPC:2} | 2 | 72
Cold cathode gauge | A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:ML | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:ML-CCG:1} |  | XF:16IDA-CT{VGC:2} | 3 | 45
Pirani gauge | A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:ML | TCG | 1 | TCG:1 | XF:16IDA-VA{Mono:ML-TCG:1} |  | XF:16IDA-CT{VGC:2} | 6 | 103
Ion pump | A | 31.5 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:ML | IP | 1 | IP:1 | XF:16IDA-VA{Mono:ML-IP:1} |  | XF:16IDA-CT{IPC:3} | 1 | 73
Cold cathode gauge | A | 31.5 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:ML | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:ML-CCG:1} |  | XF:16IDA-CT{VGC:3} | 1 | 46
Pirani gauge | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | TCG | 1 | TCG:1 | XF:16IDA-VA{Mono:Si-TCG:1} |  | XF:16IDA-CT{VGC:3} | 5 | 104
Gate valve | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | GV | 1 | GV:1 | XF:16IDA-VA{Mono:Si-GV:1}
Ion pump | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | IP | 1 | IP:1 | XF:16IDA-VA{Mono:Si-IP:1} |  | XF:16IDA-CT{IPC:3} | 2 | 70
Cold cathode gauge | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:Si-CCG:1} |  | XF:16IDA-CT{VGC:3} | 3 | 43
Pirani gauge | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | TCG | 1 | TCG:1 | XF:16IDA-VA{Mono:Si-TCG:1} |  | XF:16IDA-CT{VGC:3} | 6 | 101
Gate valve | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:Si | GV | 2 | GV:2 | XF:16IDA-VA{Mono:Si-GV:2}
Ion pump | A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN3 | IP | 1 | IP:1 | XF:16IDA-VA{Scr:SCN3-IP:1} |  | XF:16IDA-CT{IPC:4} | 1 | 75
Cold cathode gauge | A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN3 | CCG | 1 | CCG:1 | XF:16IDA-VA{Scr:SCN3-CCG:1} |  | XF:16IDA-CT{VGC:4} | 1 | 48
Pirani gauge | A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN3 | TCG | 1 | TCG:1 | XF:16IDA-VA{Scr:SCN3-TCG:1} |  | XF:16IDA-CT{VGC:4} | 5 | 105
Gate valve | A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | Scr:SCN3 | GV | 1 | GV:1 | XF:16IDA-VA{Scr:SCN3-GV:1}
Ion pump | A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | IP | 1 | IP:1 | XF:16IDA-VA{Mir:KB-IP:1} |  | XF:16IDA-CT{IPC:4} | NC 2 | 71
Cold cathode gauge | A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | CCG | 1 | CCG:1 | XF:16IDA-VA{Mir:KB-CCG:1} |  | XF:16IDA-CT{IPC:5} | NC 1 | 44
Pirani gauge | A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | TCG | 1 | TCG:1 | XF:16IDA-VA{Mir:KB-TCG:1} |  | XF:16IDA-CT{IPC:5} | NC 5 | 102
Gate valve | A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | GV | 1 | GV:1 | XF:16IDA-VA{Mir:KB-GV:1}
Ion pump | A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | Mon:SCN4 | IP | 1 | IP:1 | XF:16IDA-VA{Mon:SCN4-IP:1} |  | XF:16IDA-CT{IPC:5} | 1 | 76
Cold cathode gauge | A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | Mon:SCN4 | CCG | 1 | CCG:1 | XF:16IDA-VA{Mon:SCN4-CCG:1} |  | XF:16IDA-CT{VGC:4} | 2 | 49
Pirani gauge | A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | Mon:SCN4 | TCG | 1 | TCG:1 | XF:16IDA-VA{Mon:SCN4-TCG:1} |  | XF:16IDA-CT{VGC:4} | 6 | 106
Gauge controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 1 | VGC:1 | XF:16IDA-CT{VGC:1}
Gauge controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 2 | VGC:2 | XF:16IDA-CT{VGC:2}
Gauge controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 3 | VGC:3 | XF:16IDA-CT{VGC:3}
Gauge controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 4 | VGC:4 | XF:16IDA-CT{VGC:4}
Pump controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 1 | IPC:1 | XF:16IDA-CT{IPC:1}
Pump controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 2 | IPC:2 | XF:16IDA-CT{IPC:2}
Pump controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 3 | IPC:3 | XF:16IDA-CT{IPC:3}
Pump controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 4 | IPC:4 | XF:16IDA-CT{IPC:4}
Pump controller | A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 5 | IPC:5 | XF:16IDA-CT{IPC:5}
XFN Branch Optics (Hutch A)
XFN Branch Endstation (Hutch C)
LiX Hutch B
Ion pump | B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | Mon:SCN5 | IP | 1 | IP:1 | XF:16IDB-VA{Mon:SCN5-IP:1} |  | XF:16IDA-CT{IPC:5} | 2 | 3
Cold cathode gauge | B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | Mon:SCN5 | CCG | 1 | CCG:1 | XF:16IDB-VA{Mon:SCN5-CCG:1} |  | XF:16IDB-CT{VGC:1} | 1 | 2
Pirani gauge | B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | Mon:SCN5 | TCG | 1 | TCG:1 | XF:16IDB-VA{Mon:SCN5-TCG:1} |  | XF:16IDB-CT{VGC:1} | 5 | 24
Gate valve | B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | Mon:SCN5 | GV | 1 | GV:1 | XF:16IDB-VA{Mon:SCN5-GV:1}
Pirani gauge | B | 51 | XF | 16IDB | XF:16IDB | VA |  | VA | Chm:SS | TCG | 1 | TCG:1 | XF:16IDB-VA{Chm:SS-TCG:1} |  | XF:16IDB-CT{VGC:1} | 6 | 25
Pirani gauge | B | 51 | XF | 16IDB | XF:16IDB | VA |  | VA | Chm:SS | TCG | 2 | TCG:2 | XF:16IDB-VA{Chm:SS-TCG:2} |  | XF:16IDB-CT{VGC:2} | 1 | 38
LiX Hutch C
Gate valve | C | 51 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | GV | 1 | GV:1 | XF:16IDC-VA{Chm:SF-GV:1}
Pirani gauge | C | 58.875 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | TCG | 1 | TCG:1 | XF:16IDC-VA{Chm:SF-TCG:1} |  | XF:16IDB-CT{VGC:2} | 2 | 68
Gate valve | C | 58.9 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | GV | 2 | GV:2 | XF:16IDC-VA{Chm:SF-GV:2}
Pirani gauge | C | 59.5 | XF | 16IDC | XF:16IDC | VA |  | VA | :ES | TCG | 1 | TCG:1 | XF:16IDC-VA{:ES-TCG:1} |  | XF:16IDB-CT{VGC:2} | 3 | 69
Gate valve | C | 59.5 | XF | 16IDC | XF:16IDC | VA |  | VA | :ES | GV | 1 | GV:1 | XF:16IDC-VA{:ES-GV:1}
Pirani gauge | C | 60 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:WAXS | TCG | 1 | TCG:1 | XF:16IDC-VA{Chm:WAXS-TCG:1} |  | XF:16IDB-CT{VGC:2} | 5 | 70
Gauge controller | B | RG:D2 | XF | 16IDB | XF:16IDB | CT |  | CT |  | VGC | 1 | VGC:1 | XF:16IDB-CT{VGC:1}
Gauge controller | B | RG:D2 | XF | 16IDB | XF:16IDB | CT |  | CT |  | VGC | 2 | VGC:2 | XF:16IDB-CT{VGC:2}

## Vacuum Controllers
Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12 | Column13 | Column14 | Column15 | Column16
16 ID-A
| VAC SEC 1 |  | VAC SEC 2 |  | VAC SEC 3 |  |  |  | VAC SEC 5 |  | VAC SEC 6 |  | VAC SEC 7
| 27.4 |  | 29.1 |  | 30.7 |  |  |  | 32.7 |  | 36.3 |  | 37.7
| VGC-1A |  | VGC-2A |  | VGC-2B |  |  |  | VGC-3B |  | VGC-4A |  | *FUTURE
| 38-CCG HV |  | 41-CCG HV |  | 36-CCG HV |  |  |  | 8-CCG HV |  | 39-CCG HV |  | *109-CCG HV
| 47-CCG SIG |  | 50-CCG SIG |  | 45-CCG SIG |  |  |  | 43-CCG SIG |  | 48-CCG SIG |  | *44-CCG SIG
| 99-PIRANI |  | 107-PIRANI |  | 103-PIRANI |  |  |  | 101-PIRANI |  | 105-PIRANI |  | *102-PIRANI
| 28.2 |  |  |  | 31.5 | [7 empty] | 39
| VGC-1B |  |  |  | VGC-3A | [7 empty] | VGC-4B
| 7-CCG HV |  |  |  | 37-CCG HV | [7 empty] | 40-CCG HV
| 42-CCG SIG |  |  |  | 46-CCG SIG | [7 empty] | 49-CCG SIG
| 100-PIRANI |  |  |  | 104-PIRANI | [7 empty] | 106-PIRANI
| IPC-1 |  | IPC-2 |  | IPC-2 |  |  |  | IPC-3 |  | IPC-4 |  | IPC-4
| 74-IP |  | 77-IP |  | 72-IP |  |  |  | 70-IP |  | 75-IP |  | 71-IP
69-IP
[5 empty] | IPC-3 | [7 empty] | IPC-5
[5 empty] | IP-73 | [7 empty] | 76-IP
*FUTURE CABLES TO BE TERMINATED AND NOT CONNECTED

## Motion axis summary
Index | Component | Component ID | Full component name | Motion axis | Axis ID | Axis name | Full axis name | Motor | [9 empty] | Encoder |  |  |  |  | Switches |  |  | Control
[8 empty] | Type | Manufacturer | Model | Resolution | Current | Phase resistance (ohms) | Phase inductance (mH) | Gearbox | Microstep ratio | Overall resolution (EGU/full step) | Type | Manufacturer | Readhead Model | Interpolator Model | Resolution (micron) | Limit | Reference | Overtravel | Rack | Controller type | Controller | Axis | Motor cable number | Encoder cable number | IOC Name | EGU | PREC | MRES | ALIAS | Homing PLC
1 | LiX FOE Optics Hutch A
2 | Beam-viewing screen #1
3 | Fixed Mask
4 | White Beam Mirror | Mir:WBM | XF:16IDA-OP:LiX{Mir:WBM} | X | X | Ax:X | XF:16IDA-OP:LiX{Mir:WBM-Ax:X}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 1 | 96 | 66 | mc1 | mm | 4 | 0.0001 |  | 8
5 | White Beam Mirror | Mir:WBM | XF:16IDA-OP:LiX{Mir:WBM} | Y | Y | Ax:Y | XF:16IDA-OP:LiX{Mir:WBM-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 2 | 97 | 67 | mc1 | mm | 4 | 0.0001 |  | 9
6 | White Beam Mirror | Mir:WBM | XF:16IDA-OP:LiX{Mir:WBM} | Pitch | P | Ax:P | XF:16IDA-OP:LiX{Mir:WBM-Ax:P}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 3 | 98 | 68 | mc1 | urad | 3 | 0.011 |  | 10
7 | White Beam Mirror | Mir:WBM | XF:16IDA-OP:LiX{Mir:WBM} | Fine Pitch; Out of Scope | FP | Ax:FP | XF:16IDA-OP:LiX{Mir:WBM-Ax:FP}Mtr | [18 empty] | RG:A1 | piezo TBD |  |  | NA
8 | Beam-viewing screen #2 | Mon:SCN2 | XF:16IDA-BI:LiX{Mon:SCN2} |  | Y | Ax:Y | XF:16IDA-BI:LiX{Mon:SCN2-Ax:Y}Mtr | [18 empty] | RG:A1 |  |  |  | NA
9 | Dual Multilayer Monochromator |  |  | theta1; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 4 | 114 | 118 |  | tbd
10 | Dual Multilayer Monochromator |  |  | X1; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 5 | 115 | 119 |  | tbd
11 | Dual Multilayer Monochromator |  |  | theta2; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 6 | 116 | 120 |  | tbd
12 | Dual Multilayer Monochromator |  |  | X2; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 7 | 117 | 121 |  | tbd
13 | Dual Multilayer Monochromator |  |  | roll 2; Out of Scope | [21 empty] | RG:A1 | PicoMotor MCS |  |  | NA
14 | Dual Multilayer Monochromator |  |  | fine pitch 2; Out of Scope | [21 empty] | RG:A1 | piezoTBD |  |  | NA
15 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | theta | TH | Ax:TH | XF:16IDA-OP:LiX{Mono:Si-Ax:TH}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 1 | 78 | 51 | mc2 | mrad | 6 | 0.000109 |  | 8
16 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | X | X | Ax:X | XF:16IDA-OP:LiX{Mono:Si-Ax:X}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 2 | 79 | 52 | mc2 | mm | 3 | 0.001 |  | 9
17 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | DCM.Offset2 | OF2 | Ax:OF2 | XF:16IDA-OP:LiX{Mono:Si-Ax:OF2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 3 | 80 | 53 | mc2 | mm | 4 | 0.0002 |  | 10
18 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | DCM.Pitch2 | P2 | Ax:P2 | XF:16IDA-OP:LiX{Mono:Si-Ax:P2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 4 | 81 | 54 | mc2 | mm | 6 | 1e-06 |  | 11 | no enc
19 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | DCM.Roll2 | R2 | Ax:R2 | XF:16IDA-OP:LiX{Mono:Si-Ax:R2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 5 | 82 | 55 | mc2 | mm | 6 | 2e-06 |  | 12
20 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | DCM.finePitch piezo | FP_DCM | Ax:FP_DCM | XF:16IDA-OP:LiX{Mono:Si-Ax:FP_DCM}Mtr | [18 empty] | RG:A1 | PI? |  |  | NA
21 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | CCM.finePitch stepper | FP_CCM | Ax:FP_CCM | XF:16IDA-OP:LiX{Mono:Si-Ax:FP_CCM}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 6 | 83 | NA | mc2 | um | 4 | 0.0001 |  | 13
22 | Si(111) Monochromator (DCM/CCM) | Mono:Si | XF:16IDA-OP:LiX{Mono:Si} | DCM.finePitch readback only | FP_CCM_RDBK | Ax:FP_CCM_RDBK | XF:16IDA-OP:LiX{Mono:Si-Ax:FP_CCM_RDBK}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 7 | NA | 56 | mc2 | um | 6 | 1e-06 |  | 14 | encoder only
23 | Beam-viewing screen #3 | Scr:SCN3 | XF:16IDA-BI:LiX{Scr:SCN3} | Y | Y | Ax:Y | XF:16IDA-BI:LiX{Scr:SCN3-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 3 | 95 | NA | mc4 | um | 2 | 0.05
24 | Slits S0 | Slt:S0 | XF:16IDA-OP:LiX{Slt:S0} | X-inboard | I | Ax:I | XF:16IDA-OP:LiX{Slt:S0-Ax:I}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 5 | 91 | NA | mc4 | um | 2 | 0.05
25 | Slits S0 | Slt:S0 | XF:16IDA-OP:LiX{Slt:S0} | X-outboard | O | Ax:O | XF:16IDA-OP:LiX{Slt:S0-Ax:O}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 6 | 92 | NA | mc4 | um | 2 | 0.05
26 | Slits S0 | Slt:S0 | XF:16IDA-OP:LiX{Slt:S0} | Y-top | T | Ax:T | XF:16IDA-OP:LiX{Slt:S0-Ax:T}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 7 | 93 | NA | mc4 | um | 2 | 0.05
27 | Slits S0 | Slt:S0 | XF:16IDA-OP:LiX{Slt:S0} | Y-bottom | B | Ax:B | XF:16IDA-OP:LiX{Slt:S0-Ax:B}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 8 | 94 | NA | mc4 | um | 2 | 0.05
28 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | HFMX-upstream | HFMX-U | Ax:HFMX-U | XF:16IDA-OP:LiX{Mir:KB-Ax:HFMX-U}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 1 | 84 | 57 | mc3 | deg | 6 | 7.4485e-06
29 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | HFMX-downstream | HFMX-D | Ax:HFMX-D | XF:16IDA-OP:LiX{Mir:KB-Ax:HFMX-D}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 2 | 85 | 58 | mc3
30 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | HFMY-upstream | HFMY-U | Ax:HFMY-U | XF:16IDA-OP:LiX{Mir:KB-Ax:HFMY-U}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 3 | 86 | 59 | mc3 | urad | 2 | 0.01
31 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | HFMY-downstream | HFMY-D | Ax:HFMY-D | XF:16IDA-OP:LiX{Mir:KB-Ax:HFMY-D}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 4 | 87 | 60 | mc3 | urad | 2 | 0.01
32 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | VFMX one horizontal translation | VFMX | Ax:VFMX | XF:16IDA-OP:LiX{Mir:KB-Ax:VFMX}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 5 | 88 | 61 | mc3 | um | 3 | 0.001
33 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | VFMY-upstream | VFMY-U | Ax:VFMY-U | XF:16IDA-OP:LiX{Mir:KB-Ax:VFMY-U}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 6 | 89 | 62 | mc3 | um | 3 | 0.001
34 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | VFMY-downstream | VFMY-D | Ax:VFMY-D | XF:16IDA-OP:LiX{Mir:KB-Ax:VFMY-D}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 7 | 90 | 63 | mc3 | um | 3 | 0.001
35 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | Fine Pitch Adjustment HFM RDBK | PFH | Ax:PFH | XF:16IDA-OP:LiX{Mir:KB-Ax:PFH}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 1 | NA | 64 | mc4 | um | 3 | 0.001
36 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | Fine Pitch Adjustment HFM piezo | PFH | Ax:PFH | XF:16IDA-OP:LiX{Mir:KB-Ax:PFH}Mtr | [18 empty] | RG:A1 | PI MCS |  |  |  |  | mc | um | 3 | 0.001
37 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | Fine Pitch Adjustment VFM RDBK | PFV | Ax:PFV | XF:16IDA-OP:LiX{Mir:KB-Ax:PFV}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 2 | NA | 65 | mc4 | um | 3 | 0.001
38 | KB mirror system | Mir:KB | XF:16IDA-OP:LiX{Mir:KB} | Fine Pitch Adjustment VFM piezo | PFV | Ax:PFV | XF:16IDA-OP:LiX{Mir:KB-Ax:PFV}Mtr | [18 empty] | RG:A1 | PI MCS |  |  |  |  | mc | um | 3 | 0.001
39 | Beam-viewing screen #4 | Mon:SCN4 | XF:16IDA-OP:LiX{Mon:SCN4} | Y | Y | Ax:Y | XF:16IDA-OP:LiX{Mon:SCN4-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 4 | 108 | NA | mc4 | um | 2 | 0.05
40
41
42
43 | Secondry Source SOE Hutch B
44 | Beam Position Monitor (BPM) | Mon:BPM1 | XF:16IDB-SS:LiX{Slt:SSA} |  |  | Ax: | XF:16IDB-SS:LiX{Slt:SSA-Ax:}Mtr | [18 empty] | RG:B1 | SmartAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
45 | Secondary Source Aperture(SSA) | Slt:SSA | XF:16IDB-SS:LiX{Slt:SSA} | X | X | Ax:X | XF:16IDB-SS:LiX{Slt:SSA-Ax:X}Mtr | [18 empty] | RG:B1 | PiezoSysJena | PSJ |  |  | 0 | mcPSJ | um | 3 | 1
46 | Secondary Source Aperture(SSA) | Mon:BPM1 | XF:16IDB-SS:LiX{Slt:SSA} | Y | Y | Ax:Y | XF:16IDB-SS:LiX{Slt:SSA-Ax:Y}Mtr | [18 empty] | RG:B1 | PiezoSysJena | PSJ |  |  | 0 | mcPSJ | um | 3 | 1
47 | Attenuator | Attn | XF:16IDB-SS:LiX{Attn} | absorber set #1 | X1 | Ax:X1 | XF:16IDB-SS:LiX{Attn-Ax:X1}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 5 | B-5 | B-6 | mc5
48 | Attenuator | Attn | XF:16IDB-SS:LiX{Attn} | absorber set #2 | X2 | Ax:X2 | XF:16IDB-SS:LiX{Attn-Ax:X2}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 6 | B-9 | B-7 | mc5
49 | Attenuator | Attn | XF:16IDB-SS:LiX{Attn} | absorber set #3 | X3 | Ax:X3 | XF:16IDB-SS:LiX{Attn-Ax:X3}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 7 | B-10 | B-8 | mc5
50 | Visual Beam Monitor (VBM) | Mon:VBM | XF:16IDB-SS:LiX{Mon:VBM} | focus | F | Ax:F | XF:16IDB-SS:LiX{Mon:VBM-Ax:F}Mtr | [18 empty] | RG:B1 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
51 | Visual Beam Monitor (VBM) | Mon:VBM | XF:16IDB-SS:LiX{Mon:VBM} | zoom | Z | Ax:Z | XF:16IDB-SS:LiX{Mon:VBM-Ax:Z}Mtr | [18 empty] | RG:B1 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
52 | Shutter | Sh | XF:16IDB-SS:LiX{Sh} | open | X | Ax:X | XF:16IDB-SS:LiX{Sh-Ax:X}Mtr | [18 empty] | RG:B1 | PLC |  |  |  |  | mc | um | 3 | 1
53 | Alternative SSA | Slt:aSSA | XF:16IDB-SS:LiX{Slt:aSSA} | X | X | Ax:X | XF:16IDB-SS:LiX{Slt:aSSA-Ax:X}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 1 | B-11 |  | mc5 | um | 3 | 1
54 | Alternative SSA | Slt:aSSA | XF:16IDB-SS:LiX{Slt:aSSA} | dX | dX | Ax:dX | XF:16IDB-SS:LiX{Slt:aSSA-Ax:dX}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 2 | B-12 |  | mc5
55 | Alternative SSA | Slt:aSSA | XF:16IDB-SS:LiX{Slt:aSSA} | Y | Y | Ax:Y | XF:16IDB-SS:LiX{Slt:aSSA-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 3 | B-22 |  | mc5
56 | Alternative SSA | Slt:aSSA | XF:16IDB-SS:LiX{Slt:aSSA} | dY | dY | Ax:dY | XF:16IDB-SS:LiX{Slt:aSSA-Ax:dY}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 4 | B-23 |  | mc5
57
58
59 | Experimental End Station Enclosure EESE Hutch C
60 | Secondary focusing working distance | Stg:WD | XF:16IDC-SF:LiX{Stg:WD} | Z | Z | Ax:Z | XF:16IDC-SF:LiX{Stg:WD-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 1 | C-7 |  | mc9 | um | 3 | 1
61 | Secondary focusing working distance | Stg:WD | XF:16IDC-SF:LiX{Stg:WD} | Z1 | Z1 | Ax:Z1 | XF:16IDC-SF:LiX{Stg:WD-Ax:Z1}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 2 | C-8 |  | mc9 | um | 3 | 1
62 | Harmonic Regection Mirror HRM1 | Mir:HRM1 | XF:16IDC-SF:LiX{Mir:HRM1} | Y | Y | Ax:Y | XF:16IDC-SF:LiX{Mir:HRM1-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 2 | C-20 |  | mc6 | um | 3 | 1
63 | Harmonic Regection Mirror HRM1 | Mir:HRM1 | XF:16IDC-SF:LiX{Mir:HRM1} | theta | TH | Ax:TH | XF:16IDC-SF:LiX{Mir:HRM1-Ax:TH}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 1 | C-19 |  | mc6 | um | 3 | 1
64 | Transfocator CRLs | Trf | XF:16IDC-SF:LiX{Trf} | X Upstream | XU | Ax:XU | XF:16IDC-SF:LiX{Trf-Ax:XU}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 5 | C-11 |  | mc6 | um | 3 | 1
65 | Transfocator CRLs | Trf | XF:16IDC-SF:LiX{Trf} | Y Upstream | YU | Ax:YU | XF:16IDC-SF:LiX{Trf-Ax:YU}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 6 | C-12 |  | mc6 | um | 3 | 1
66 | Transfocator CRLs | Trf | XF:16IDC-SF:LiX{Trf} | X Downstream | XD | Ax:XD | XF:16IDC-SF:LiX{Trf-Ax:XD}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 7 | C-13 |  | mc6 | um | 3 | 1
67 | Transfocator CRLs | Trf | XF:16IDC-SF:LiX{Trf} | Y Downstream | YD | Ax:YD | XF:16IDC-SF:LiX{Trf-Ax:YD}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 8 | C-14 |  | mc6 | um | 3 | 1
68 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-SF:LiX{Mir:HRM2} | Y | Y | Ax:Y | XF:16IDC-SF:LiX{Mir:HRM2-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 3 | C-21 |  | mc6 | um | 3 | 1
69 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-SF:LiX{Mir:HRM2} | theta, tilt | TH | Ax:TH | XF:16IDC-SF:LiX{Mir:HRM2-Ax:TH}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 8 | C-22 | C-23 | mc5 | rad | 3 | 1
70 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-SF:LiX{Mir:HRM2} | Upstream Bend; with LVDT | BU | Ax:BU | XF:16IDC-SF:LiX{Mir:HRM2-Ax:BU}Mtr | [18 empty] | RG:C2 | Picomotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
71 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-SF:LiX{Mir:HRM2} | Downstream Bend; With LVDT | BD | Ax:BD | XF:16IDC-SF:LiX{Mir:HRM2-Ax:BD}Mtr | [18 empty] | RG:C2 | Picomotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
72 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-SF:LiX{Slt:DDA} | X | X | Ax:X | XF:16IDC-SF:LiX{Slt:DDA-Ax:X}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 1 | C-15 |  | mc7 | um | 3 | 1
73 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-SF:LiX{Slt:DDA} | Y | Y | Ax:Y | XF:16IDC-SF:LiX{Slt:DDA-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 2 | C-16 |  | mc7 | um | 3 | 1
74 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-SF:LiX{Slt:DDA} | dX | dX | Ax:dX | XF:16IDC-SF:LiX{Slt:DDA-Ax:dX}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 3 | C-17 |  | mc7 | um | 3 | 1
75 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-SF:LiX{Slt:DDA} | dY | dY | Ax:dY | XF:16IDC-SF:LiX{Slt:DDA-Ax:dY}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 4 | C-18 |  | mc7 | um | 3 | 1
76 | Beam Position Monitor (BPM) | Mon:BPM2 | XF:16IDC-SF:LiX{Mon:BPM2} | Y | Y | Ax:Y | XF:16IDC-SF:LiX{Mon:BPM2-Ax:Y}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
77 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-SF:LiX{Slt:G1} | X | X | Ax:X | XF:16IDC-SF:LiX{Slt:G1-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 1 | C-9 |  | mc10 | um | 3 | 1
78 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-SF:LiX{Slt:G1} | dX | dX | Ax:dX | XF:16IDC-SF:LiX{Slt:G1-Ax:dX}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 2 | C-10 |  | mc10 | um | 3 | 1
79 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-SF:LiX{Slt:G1} | Y | Y | Ax:Y | XF:16IDC-SF:LiX{Slt:G1-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 3 | C-25 |  | mc10 | um | 3 | 1
80 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-SF:LiX{Slt:G1} | dY | dY | Ax:dY | XF:16IDC-SF:LiX{Slt:G1-Ax:dY}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 4 | C-26 |  | mc10 | um | 3 | 1
81 | In-air transmission measurement module
82 | Coarse-resolution scanning | Stg:CR | XF:16IDC-ES:LiX:InAir{Stg:CR} | scanX | scanX | Ax:scanX | XF:16IDC-ES:LiX:InAir{Stg:CR-Ax:scanX}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 1 | C-34 |  | mc8 | um | 3 | 1
83 | Coarse-resolution scanning | Stg:CR | XF:16IDC-ES:LiX:InAir{Stg:CR} | scanY | scanY | Ax:scanY | XF:16IDC-ES:LiX:InAir{Stg:CR-Ax:scanY}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 2 | C-35 |  | mc8 | um | 3 | 1
84 | Fine-resolution scanning | Stg:FR | XF:16IDC-ES:LiX:InAir{Stg:FR} | scanX | scanX | Ax:scanX | XF:16IDC-ES:LiX:InAir{Stg:FR-Ax:scanX}Mtr | [18 empty] | RG:C2 | PI MCS | PI |  |  |  | mcPI | um | 3 | 1
85 | Fine-resolution scanning | Stg:FR | XF:16IDC-ES:LiX:InAir{Stg:FR} | scanY | scanY | Ax:scanY | XF:16IDC-ES:LiX:InAir{Stg:FR-Ax:scanY}Mtr | [18 empty] | RG:C2 | PI MCS | PI |  |  |  | mcPI | um | 3 | 1
86 | Fine-resolution scanning | Stg:FR | XF:16IDC-ES:LiX:InAir{Stg:FR} | azimuthal | Rot | Ax:Rot | XF:16IDC-ES:LiX:InAir{Stg:FR-Ax:Rot}Mtr | [18 empty] | RG:C2 | GBLV | 6 | 4 | C-24 |  | mc6 | um | 3 | 1
87 | Microscope | Mscp | XF:16IDC-ES:LiX:InAir{Mscp} | X | X | Ax:X | XF:16IDC-ES:LiX:InAir{Mscp-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 4 | C-31 |  | mc8 | um | 3 | 1
88 | Microscope | Mscp | XF:16IDC-ES:LiX:InAir{Mscp} | Y | Y | Ax:Y | XF:16IDC-ES:LiX:InAir{Mscp-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 5 | C-32 |  | mc8 | um | 3 | 1
89 | Microscope | Mscp | XF:16IDC-ES:LiX:InAir{Mscp} | focus | F | Ax:F | XF:16IDC-ES:LiX:InAir{Mscp-Ax:F}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 8 | C-33 |  | mc10 | um | 3 | 1
90 | Microscope | Mscp | XF:16IDC-ES:LiX:InAir{Mscp} | polarizer/analyze | P | Ax:P | XF:16IDC-ES:LiX:InAir{Mscp-Ax:P }Mtr | [18 empty] | RG:B1 | GBLV | 7 | 7 | C-106 |  | mc7 | um | 3 | 1
91 | Microscope | Mscp | XF:16IDC-ES:LiX:InAir{Mscp} | zoom | Z | Ax:Z | XF:16IDC-ES:LiX:InAir{Mscp-Ax:Z}Mtr | [19 empty] | Fetura RS-232
92 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:LiX:InAir{Slt:G2} | X | X | Ax:X | XF:16IDC-ES:LiX:InAir{Slt:G2-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 1 | C-27 |  | mc11 | um | 3 | 1
93 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:LiX:InAir{Slt:G2} | dX | dX | Ax:dX | XF:16IDC-ES:LiX:InAir{Slt:G2-Ax:dX}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 2 | C-28 |  | mc11 | um | 3 | 1
94 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:LiX:InAir{Slt:G2} | Y | Y | Ax:Y | XF:16IDC-ES:LiX:InAir{Slt:G2-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 3 | C-29 |  | mc11 | um | 3 | 1
95 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:LiX:InAir{Slt:G2} | dY | dY | Ax:dY | XF:16IDC-ES:LiX:InAir{Slt:G2-Ax:dY}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 4 | C-30 |  | mc11 | um | 3 | 1
96 | In-vaccum GISAXS/GID module
97 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | X | X | Ax:X | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:X}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
98 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | Y | Y | Ax:Y | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:Y}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
99 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | Z | Z | Ax:Z | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:Z}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
100 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | U | U | Ax:U | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:U}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
101 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | V | V | Ax:V | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:V}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
102 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | W | W | Ax:W | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:W}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
103 | Hexapod | Stg:Hex | XF:16IDC-ES:LiX:GI{Stg:Hex} | phi | phi | Ax:phi | XF:16IDC-ES:LiX:GI{Stg:Hex-Ax:phi}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
104 | Guard Slits 2 in VAC | Slt:G2 | XF:16IDC-ES:LiX:GI{Slt:G2} | X | X | Ax:X | XF:16IDC-ES:LiX:GI{Slt:G2-Ax:X}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
105 | Guard Slits 2 in VAC | Slt:G2 | XF:16IDC-ES:LiX:GI{Slt:G2} | dX | dX | Ax:dX | XF:16IDC-ES:LiX:GI{Slt:G2-Ax:dX}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
106 | Sample handling
107 | Solution sample handler | Rbt:Sol | XF:16IDC-ES:LiX{Rbt:Sol} | X | X | Ax:X | XF:16IDC-ES:LiX{Rbt:Sol-Ax:X}Mtr | [18 empty] | RG:C2 | TBD MCS/GBLV | TBD |  |  |  | mcTBD | um | 3 | 1
108 | Solution sample handler | Rbt:Sol | XF:16IDC-ES:LiX{Rbt:Sol} | Y | Y | Ax:Y | XF:16IDC-ES:LiX{Rbt:Sol-Ax:Y}Mtr | [18 empty] | RG:C2 | TBD MCS/GBLV | TBD |  |  |  | mcTBD | um | 3 | 1
109 | Staubli TX40 robot | Rbt:Staubli | XF:16IDC-ES:LiX{Rbt:Staubli} | routins will be called via PVs | TBD | Ax:TBD | XF:16IDC-ES:LiX{Rbt:Staubli-Ax:TBD }Mtr | [18 empty] | RG:C2 | Staubli MCS | Staubli |  |  |  | mcStaubli | um | 3 | 1
110 | Detector system
111 | WAXS detector 1 | Det:WAXS1 | XF:16IDC-DT:LiX{Det:WAXS1} | X | X | Ax:X | XF:16IDC-DT:LiX{Det:WAXS1-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 3 | C-36 |  | mc9 | um | 3 | 1
112 | WAXS detector 1 | Det:WAXS1 | XF:16IDC-DT:LiX{Det:WAXS1} | Y | Y | Ax:Y | XF:16IDC-DT:LiX{Det:WAXS1-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 4 | C-37 |  | mc9 | um | 3 | 1
113 | WAXS detector 1 | Det:WAXS1 | XF:16IDC-DT:LiX{Det:WAXS1} | Z | Z | Ax:Z | XF:16IDC-DT:LiX{Det:WAXS1-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 5 | C-38 |  | mc9 | um | 3 | 1
114 | WAXS detector 2 | Det:WAXS2 | XF:16IDC-DT:LiX{Det:WAXS2} | X | X | Ax:X | XF:16IDC-DT:LiX{Det:WAXS2-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 6 | C-39 |  | mc9 | um | 3 | 1
115 | WAXS detector 2 | Det:WAXS2 | XF:16IDC-DT:LiX{Det:WAXS2} | Y | Y | Ax:Y | XF:16IDC-DT:LiX{Det:WAXS2-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 7 | C-40 |  | mc9 | um | 3 | 1
116 | WAXS detector 2 | Det:WAXS2 | XF:16IDC-DT:LiX{Det:WAXS2} | Z | Z | Ax:Z | XF:16IDC-DT:LiX{Det:WAXS2-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 8 | C-41 |  | mc9 | um | 3 | 1
117 | WAXS beamstop | BS:WAXS | XF:16IDC-DT:LiX{BS:WAXS} | X | X | Ax:X | XF:16IDC-DT:LiX{BS:WAXS-Ax:X}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
118 | WAXS beamstop | BS:WAXS | XF:16IDC-DT:LiX{BS:WAXS} | Y | Y | Ax:Y | XF:16IDC-DT:LiX{BS:WAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
119 | WAXS beamstop | BS:WAXS | XF:16IDC-DT:LiX{BS:WAXS} | Z | Z | Ax:Z | XF:16IDC-DT:LiX{BS:WAXS-Ax:Z}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
120 | SAXS detector | Det:SAXS | XF:16IDC-DT:LiX{Det:SAXS} | X | X | Ax:X | XF:16IDC-DT:LiX{Det:SAXS-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 6 | C-48 |  | mc8 | um | 3 | 1
121 | SAXS detector | Det:SAXS | XF:16IDC-DT:LiX{Det:SAXS} | Y | Y | Ax:Y | XF:16IDC-DT:LiX{Det:SAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 7 | C-49 |  | mc8 | um | 3 | 1
122 | SAXS detector (FlightPath Z) | Det:SAXS | XF:16IDC-DT:LiX{Det:SAXS} | Z | Z | Ax:Z | XF:16IDC-DT:LiX{Det:SAXS-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 8 | C-50 | C-45 C-46 | mc8 | um | 3 | 1
123 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | X | X | Ax:X | XF:16IDC-DT:LiX{BS:SAXS-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 5 | C-51 |  | mc11 | um | 3 | 1
124 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | Y | Y | Ax:Y | XF:16IDC-DT:LiX{BS:SAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 6 | C-52 |  | mc11 | um | 3 | 1
125 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | R1 | R1 | Ax:R1 | XF:16IDC-DT:LiX{BS:SAXS-Ax:R1}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 5 | C-65 |  | mc10 | um | 3 | 1
126 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | R2 | R2 | Ax:R2 | XF:16IDC-DT:LiX{BS:SAXS-Ax:R2}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 6 | C-66 |  | mc10 | um | 3 | 1
127 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | R3 | R3 | Ax:R3 | XF:16IDC-DT:LiX{BS:SAXS-Ax:R3}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 7 | C-67 |  | mc10 | um | 3 | 1
128 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | tilt1 | T1 | Ax:T1 | XF:16IDC-DT:LiX{BS:SAXS-Ax:T1}Mtr | [18 empty] | RG:C2 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
129 | SAXS beam stop | BS:SAXS | XF:16IDC-DT:LiX{BS:SAXS} | tilt2 | T2 | Ax:T2 | XF:16IDC-DT:LiX{BS:SAXS-Ax:T2}Mtr | [18 empty] | RG:C2 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
Note on B98: This is part of the GISAXS/GID module

## GBLV Motor Controllers
Beamline: | LiX
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
Controler | Controller #1 | Controller #2 | Controller #3 | Controller #4 | Controller #5 | Controller #6 | Controller #7 | Controller #8 | Controller #9 | Controller #10 | Controller #11 | Controller #12 | Controller #13
Location | Rack A-1 | Rack A-1 | Rack A-1 | Rack A-1 | Rack B-1 | Rack B-1 | Rack B-1 | Rack C-2 | Rack C-2 | Rack C-2 | Rack C-2 | Spare | Spare
Model # | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000
Channels' description | 8x5A | 8x5A | 8x5A | 8x5A | 4x1A, 4x0.25A | 8x5A | 4x1A, 4x0.25A | 8x5A | 8x5A | 4x1A, 4x0.25A | 4x1A, 4x0.25A | 8x5A | 4x1A, 4x0.25A
Channel 1 | White Beam Mirror X | Monochromator theta | KB mirror sys HFMX-upstream | KB HFM fine pitch (readback only) | Alternative SSA X (450mA) | HRM1 theta | Divergence-defining aperture X (450mA) | In-air measurement Scan X | Secondary Focusing working distance Z | Slits Guard SG1 X (450mA) | Slits Guard SG2 X (450mA)
Channel 2 | White beam mirror Y | Monochromator X | KB mirror sys HFMX-downstream | KB VFM fine pitch (readback only) | Alternative SSA dX (450mA) | HRM1 Y | Divergence-defining aperture Y(450mA) | In-air measurement Scan Y | Secondary Focusing working distance Z1 | Slits Guard SG1 dX (450mA) | Slits Guard SG2 dX (450mA)
Channel 3 | White beam mirror Pitch | Monochromator DCM.Offset2 | KB mirror sys HFMY-upstream | Beam viewing screen #3 Y | Alternative SSA Y (450mA) | HRM2 Y | Divergence-defining aperture dX (450mA) |  | WAXS detector1 X | Slits Guard SG1 Y (450mA) | Slits Guard SG2 Y (450mA)
Channel 4 | Multilayer monochromator theta1 | Monochromator DCM.Pitch2 | KB mirror sys HFMY-downstream | Beam viewing screen #4 Y | Alternative SSA dY (450mA) | Fine-resolution scanning (3000 mA) | Divergence-defining aperture dY (450mA) | microscope X | WAXS detector1 Y | Slits Guard SG1 dY (450mA) | Slits Guard SG2 dY (450mA)
Channel 5 | Multilayer monochromator X1 | Monochromator DCM.Roll2 | KB mirror sys VFMX | Slit S0 X-inboard | Attenuator X1 (75mA) | Transfocator X-upstream |  | microscope Y | WAXS detector1 Z | SAXS beam stop R1 (130mA) | SAXS beam stop X (170mA)
Channel 6 | Multilayer monochromator theta2 | Monochromator CCM.finePitch | KB mirror sys VFMY-upstream | Slit S0 X-outboard | Attenuator X2 (75mA) | Transfocator Y-upstream |  | SAXS detector X | WAXS detector2 X | SAXS beam stop R2 (130mA) | SAXS beam stop Y (170mA)
Channel 7 | Multilayer monochromator X2 | Monochromator CCM.finePitch (readback only) ?? | KB mirror sys VFMY-downstream | Slit S0 Y-top | Attenuator X3 (75mA) | Transfocator X-downstream | microscope polarizer/analyzer (130 mA) | SAXS detector Y | WAXS detector2 Y | SAXS beam stop R3 (130mA)
Channel 8 |  |  |  | Slit S0 Y-bottom | HRM2 theta | Transfocator Y-downstream |  | SAXS detector Z | WAXS detector2 Z | parfocus (250 mA)
Note on D12: VFM will now have only one horizontal translation

## Motor Channels
BEAMLINE: | Lix
| Totals → | 63 | 9 | 16 | >> This Subtotal does NOT include in the Front End. <<
Device | Motion Axis | Stepper Motor | Piezo | Other | Maximum | Minimum | Units | Notes | 5A amplifier | 0.1A-1A amplifier | 25mA-250mA
Front End
XBPM1 | X | 1 | [5 empty] | , ????
| Y | 1 | [5 empty] | , ????
XY Slits | X-inboard | 1 | [5 empty] | , ????
| X-outboard | 1 | [5 empty] | , ????
| Y-top | 1 | [5 empty] | , ????
| Y-bottom | 1 | [5 empty] | , ????
| Subtotal | 6 | 0 |  |  |  |  | >> This Subtotal NOT included in the Total above. <<
Hutch A (FOE)
Beam-viewing screen #1 | [7 empty] | Pneumatic
White beam mirror | X | 1 | [5 empty] | Inboard and Outboard, ~2A | 1
| Y | 1 | [5 empty] | Elevation, ~2A | 1
| Pitch | 1 | [5 empty] | ~2A | 1
| finePitch |  | 1
Beam-viewing screen #2 | Y | [6 empty] | Pneumatic
Multilayer monochromator | theta1 | 1 | [5 empty] | ~2A | 1
| X1 | 1 | [5 empty] | ~2A | 1
| theta2 | 1 | [5 empty] | ~2A | 1
| X2 | 1 | [5 empty] | ~2A | 1
| roll2 |  |  | 1 |  |  |  | likely something like PicoMotor
| finePitch2 |  | 1
Si(111) monochromator (DCM and CCM) | theta | 1 | [5 empty] | Bragg angle, ~2A | 1
| X | 1 | [5 empty] | switching between the DCM and the CCM, ~2A | 1
| DCM.Offset2 | 1 | [5 empty] | distance bwtween the two crystals, ~1A |  | 1
| DCM.Pitch2 | 1 | [5 empty] | second crystal pitch, ~1A |  | 1
| DCM.Roll2 | 1 | [5 empty] | second crystal roll, ~1A |  | 1
| DCM.finePitch2 |  | 1 |  |  |  |  | second crystal fine pitch, piezo
| CCM.finePitch | 1 | [5 empty] | ~1A |  | 1
Beam viewing screen #3 | Y | 1 | [5 empty] | ~2A | 1
Slits (S0) | X-inboard | 1 | [5 empty] | ~2A | 1
| X-outboard | 1 | [5 empty] | ~2A | 1
| Y-top | 1 | [5 empty] | ~2A | 1
| Y-bottom | 1 | [5 empty] | ~2A | 1
KB mirror system | HFMX-upstream | 1 | [5 empty] | 2.5A | 1
| HFMX-downstream | 1 | [5 empty] | 2.5A | 1
| HFMY-upstream | 1 | [5 empty] | 2.5A | 1
| HFMY-downstream | 1 | [5 empty] | 2.5A | 1
| VFMX | 1 | [5 empty] | 2.5A | 1
[8 empty] | VFMX-downstream has been removed | 1
| VFMY-upstream | 1 | [5 empty] | 2.5A
| VFMY-downstream | 1 | [5 empty] | 2.5A | 1
| fine pitch adjustments |  | 2 |  |  |  |  | PI piezo actuator, controller to be provided by us
Beam viewing screen #4 | Y | 1 | [5 empty] | ~2A | 1
| Subtotal | 26 | 5 | 1 |  |  |  | Total axis: | 26 |  | 0
Hutch B (SOE)
BPM | X |  |  | 1 |  |  |  | SmarAct motor
SSA | dX |  | 1 |  |  |  |  | horizontal opening, PiezoSystem Jena
| dY |  | 1 |  |  |  |  | vertical opening, PiezoSystem Jena
Attenuator | X1, X2, X3 | 3 | [5 empty] | low-current stepper, 75mA |  |  | 3
Video beam monitor | focus |  |  | 1 |  |  |  | open-loop PicoMotor,
| zoom |  |  | 1 |  |  |  | open-loop PicoMotor, not essential
Shutter/Scintilator | theta |  |  | 1 |  |  |  | Bistable solenoid, have driver, requires 1 TTL input (PLC?), also need to coordinate with detector trigger (software or TTL)
Alternative SSA | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
| Subtotal | 7 | 2 | 4 |  |  |  | Total axis | 0 | 4 | 3
Hutch C (Experimental Station)
Secondary Focusing working distance | Z | 1 | [5 empty] | Adjust working distance of the transfocator in vacuum, ~2A | 1
| Z1 | 1 | [5 empty] | Adjust the location of the SF chamber, ~2A | 1
Transfocator | X-upstram | 1 | [5 empty] | Kohzu AXA04A-R101-1F Motor PKP214D06B 2ph .6A | 1
| Y-upstream | 1 | [5 empty] | Kohzu AZA05A-W1-1C Motor: PKP214D06B 2ph .6A | 1
| X-downstream | 1 | [5 empty] | ~1.5A | 1
| Y-downstream | 1 | [5 empty] | ~1.5A | 1
| CRLgroup1-9 |  |  | 9 |  |  |  | brushless DC motors, each motor require two inputs (enable and direction) from the PLC
Divergence-defining aperture (DDA) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
HRM1 | theta | 1 | [5 empty] | Kohzu SA04B-RT-2N Motor: PKP223D15B 2ph, 1.5A | 1
| Y | 1 | [5 empty] | Kohzu; ZA07A-X1-1T Motor: PKP243D15B,2ph 1.5A | 1
HRM2 | theta, tilt | 1 | [5 empty] | Tilt: Newport URS50BPPV6 15 deg/step, 0.16A |  |  | 1
| Y | 1 | [5 empty] | Kohzu; ZA07A-X1-1T Motor: PKP243D15B,2ph 1.5A | 1
Guard slits 1 (Sg1) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
BIM1 | Y |  |  | 1 |  |  |  | SmarAct, moves a diamond quad-BPM
Guard slits 2 (Sg2 for grazing incidence) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
Guard slits 2 (Sg2 for transmission) | dX, dY |  |  | 2 |  |  |  | SmarAct
| Subtotal | 22 | 0 | 12
Microscope | zoom, focus | 2 | [5 empty] | , ????
| polarizer/analyzer | 1 | [5 empty] | , ????
| alignment X and Y | 2 | [5 empty] | , ????
| position X , Y, focus | 3 | [5 empty] | , ????
In-air measurement | scanX, scanY | 2 | [5 empty] | scanX, scanY, stepper 670 mA/phase |  | 2
| finescanX, finescanY |  | 2 |  |  |  |  | pizeo scanning stage from PI
| azimuthal | 1 | [5 empty] | stepper 2 phase, 1.8 deg/step NEMA23 ~1A | 1
in-vacuum GISAXS/GID with hexapod | X, Y, Z, U, V, W, phi |  |  | 7 |  |  |  | already have controller from miCos/PI
GISAXS/GID | X,Y | 2 | [5 empty] | stepper, probably size 23, ~1A | 2
| tiltX, tiltY |  |  | 2 |  |  |  | Newport
WAXS detectors | X1, Y1, X1, X2, Y2, Z2 | 6 | [5 empty] | ~2.5A | 6
WAXS beamstop | X, Y, Z | 3 | [5 empty] | SmarAct
SAXS detector | X, Y, Z | 3 | [5 empty] | large steppers | 3
SAXS beam stop | R1, R2, R3 | 3 | [5 empty] | low-current steppers, ~130mA |  |  | 3
| X, Y | 2 | [5 empty] | small servo motors, 170mA |  |  | 2
| tilt1, tilt2 |  |  | 2 |  |  |  | Open-loop PicoMotors
| Subtotal | 30 | 2 | 11 |  |  |  | Total minus ????: | 21 | 14 | 6
| total | 85 | 9 | 28 | 122

## 2015FebRackLayout
These are rackes on the roofs
| RG A-1 |  |  | RG-A1 back |  |  | RG-A2 front |  |  |  | RG-B1 |  |  |  | RG-C1 |  |  | RG-C2
42 |  | 42 | 42 |  | 42 | 42 |  |  | 42 | 42 |  |  | 42 | 42 |  | 42 | 42 |  | 42
41 |  | 41 | 41 | fiber runs to this switch | 41 | 41 |  |  | 41 | 41 |  |  | 41 | 41 |  | 41 | 41 |  | 41
40 |  | 40 | 40 | switch | 40 | 40 | TEMP. CONTROL |  | 40 | 40 |  |  | 40 | 40 | TEMP. CONTROL | 40 | 40 | TEMP. CONTROL | 40
39 |  | 39 | 39 | cable management | 39 | 39 | SPACER |  | 39 | 39 | MOXA |  | 39 | 39 | fiber patch pannel | 39 | 39 | fiber patch pannel | 39
38 | IOC | 38 | 38 |  | 38 | 38 | MOXA |  | 38 | 38 | SPACER |  | 38 | 38 | switch x48ports | 38 | 38 | MOXA | 38
37 | MOXA | 37 | 37 |  | 37 | 37 | SPACER |  | 37 | 37 | UPS |  | 37 | 37 | cable management | 37 | 37 |  | 37
36 | SPACER | 36 | 36 |  | 36 | 36 | GAUGE CONTROL | GAUGE CONTROL | 36 | 36 |  |  | 36 | 36 | cable management | 36 | 36 |  | 36
35 | UPS | 35 | 35 |  | 35 | 35 |  |  | 35 | 35 | SPACER |  | 35 | 35 | reserve switch future | 35 | 35 |  | 35
34 |  | 34 | 34 |  | 34 | 34 | SPACER |  | 34 | 34 | beam stabilization |  | 34 | 34 | cable managent future | 34 | 34 |  | 34
33 | SPACER | 33 | 33 |  | 33 | 33 | GAUGE CONTROL | GAUGE CONTROL | 33 | 33 | TEMP. CONTROL |  | 33 | 33 | cable managent future | 33 | 33 | Archiver IOC | 33
32 | Bimorph HV | 32 | 32 |  | 32 | 32 |  |  | 32 | 32 |  |  | 32 | 32 | IOC | 32 | 32 |  | 32 | ioc should be 32 or higher because of the ppwer strips
31 |  | 31 | 31 |  | 31 | 31 | SPACER |  | 31 | 31 |  |  | 31 | 31 | from B/C hutch patch pannel x12 | 31 | 31 | BNC/lemo patch panel x16 | 31
30 |  | 30 | 30 |  | 30 | 30 | VACUUM PLC |  | 30 | 30 |  |  | 30 | 30 | network patch pannel x24 | 30 | 30 | TIMING SYSTEM VME | 30 |  | Inside B hutch
29 | Bimorph HV | 29 | 29 |  | 29 | 29 |  |  | 29 | 29 |  |  | 29 | 29 | SPACER | 29 | 29 |  | 29 |  | BNC/lemo patch panel x8
28 |  | 28 | 28 |  | 28 | 28 |  |  | 28 | 28 |  |  | 28 | 28 | UPS | 28 | 28 |  | 28
27 |  | 27 | 27 |  | 27 | 27 |  |  | 27 | 27 |  |  | 27 | 27 |  | 27 | 27 |  | 27
26 | FAN | 26 | 26 |  | 26 | 26 | SPACER |  | 26 | 26 | GAUGE CONTROL | GAUGE CONTROL | 26 | 26 | SPACER | 26 | 26 |  | 26
25 | TEMP. CONTROL | 25 | 25 |  | 25 | 25 | ION PUMP CONTROL |  | 25 | 25 |  |  | 25 | 25 | PPU | 25 | 25 |  | 25
24 | Bimorph Controller | 24 | 24 |  | 24 | 24 |  |  | 24 | 24 | SPACER |  | 24 | 24 | P1M control | 24 | 24 | VME shelf | 24
23 |  | 23 | 23 |  | 23 | 23 |  |  | 23 | 23 | EPS BEAMLINE PLC |  | 23 | 23 | SPACER | 23 | 23 | SPACER | 23 |  | Inside C hutch
22 | SPACER | 22 | 22 |  | 22 | 22 | SPACER |  | 22 | 22 |  |  | 22 | 22 | P300K control | 22 | 22 | DELTA-TAU PSU | 22 |  | BNC/lemo patch panel x8
21 | DELTA-TAU PSU | 21 | 21 |  | 21 | 21 | ION PUMP CONTROL |  | 21 | 21 |  |  | 21 | 21 | P300K control | 21 | 21 | SPACER | 21
20 | SPACER | 20 | 20 |  | 20 | 20 |  |  | 20 | 20 |  |  | 20 | 20 |  | 20 | 20 | DELTA-TAU#11 low current | 20
19 | DELTA-TAU#4 | 19 | 19 |  | 19 | 19 |  |  | 19 | 19 | SPACER |  | 19 | 19 |  | 19 | 19 |  | 19
18 |  | 18 | 18 |  | 18 | 18 | SPACER |  | 18 | 18 | DELTA-TAU PSU |  | 18 | 18 |  | 18 | 18 |  | 18
17 |  | 17 | 17 |  | 17 | 17 | ION PUMP CONTROL |  | 17 | 17 | SPACER |  | 17 | 17 |  | 17 | 17 | SPACER | 17
16 | SPACER | 16 | 16 |  | 16 | 16 |  |  | 16 | 16 | DELTA-TAU#7 low current |  | 16 | 16 |  | 16 | 16 | DELTA-TAU#10 low current | 16
15 | DELTA-TAU#3 | 15 | 15 |  | 15 | 15 |  |  | 15 | 15 |  |  | 15 | 15 |  | 15 | 15 |  | 15
14 |  | 14 | 14 |  | 14 | 14 | SPACER |  | 14 | 14 |  |  | 14 | 14 |  | 14 | 14 |  | 14
13 |  | 13 | 13 |  | 13 | 13 | ION PUMP CONTROL |  | 13 | 13 | SPACER |  | 13 | 13 |  | 13 | 13 | SPACER | 13
12 | SPACER | 12 | 12 |  | 12 | 12 |  |  | 12 | 12 | DELTA-TAU#6 |  | 12 | 12 | Chillers for Pilatus detectors | 12 | 12 | DELTA-TAU#9 | 12
11 | DELTA-TAU#2 | 11 | 11 |  | 11 | 11 |  |  | 11 | 11 |  |  | 11 | 11 |  | 11 | 11 |  | 11
10 |  | 10 | 10 |  | 10 | 10 | SPACER |  | 10 | 10 |  |  | 10 | 10 |  | 10 | 10 |  | 10
9 |  | 9 | 9 |  | 9 | 9 | ION PUMP CONTROL |  | 9 | 9 | SPACER |  | 9 | 9 |  | 9 | 9 | SPACER | 9
8 | SPACER | 8 | 8 |  | 8 | 8 |  |  | 8 | 8 | DELTA-TAU#5 low current |  | 8 | 8 |  | 8 | 8 | DELTA-TAU#8 | 8
7 | DELTA-TAU#1 | 7 | 7 |  | 7 | 7 |  |  | 7 | 7 |  |  | 7 | 7 |  | 7 | 7 |  | 7
6 |  | 6 | 6 |  | 6 | 6 | SPACER |  | 6 | 6 |  |  | 6 | 6 |  | 6 | 6 |  | 6 |  | Inside B hutch
5 |  | 5 | 5 |  | 5 | 5 | RGA controller |  | 5 | 5 | network patch pannel x12 |  | 5 | 5 |  | 5 | 5 | network patch pannel x12 | 5 |  | network patch pannel x4
4 |  | 4 | 4 |  | 4 | 4 |  |  | 4 | 4 |  |  | 4 | 4 | water tray for leaks | 4 | 4 |  | 4
3 |  | 3 | 3 |  | 3 | 3 |  |  | 3 | 3 |  |  | 3 | 3 |  | 3 | 3 |  | 3 |  | Inside C hutch
2 |  | 2 | 2 |  | 2 | 2 |  |  | 2 | 2 |  |  | 2 | 2 |  | 2 | 2 |  | 2 |  | network patch pannel x8
1 |  | 1 | 1 |  | 1 | 1 |  |  | 1 | 1 |  |  | 1 | 1 |  | 1 | 1 |  | 1
controller for WAXS positioning stages
UPS ~4kW
Note on Q7: SmarAct controller ??, pico motor controller?? BPM electronics??

## 2015FebRackLayout_Floor
Thise are racks on the floor.
| RG Floor-1 |  |  | RG-Floor2
42 |  | 42 | 42 |  | 42
41 |  | 41 | 41 |  | 41
40 | fiber patch pannel | 40 | 40 |  | 40
39 |  | 39 | 39 |  | 39
38 |  | 38 | 38 |  | 38
37 |  | 37 | 37 |  | 37
36 |  | 36 | 36 |  | 36
35 |  | 35 | 35 |  | 35
34 |  | 34 | 34 |  | 34
33 |  | 33 | 33 |  | 33
32 |  | 32 | 32 |  | 32
31 |  | 31 | 31 |  | 31
30 |  | 30 | 30 |  | 30
29 |  | 29 | 29 |  | 29
28 |  | 28 | 28 |  | 28
27 |  | 27 | 27 |  | 27
26 |  | 26 | 26 |  | 26
25 |  | 25 | 25 |  | 25
24 |  | 24 | 24 |  | 24
23 |  | 23 | 23 |  | 23
22 |  | 22 | 22 |  | 22
21 |  | 21 | 21 |  | 21
20 |  | 20 | 20 |  | 20
19 |  | 19 | 19 |  | 19
18 |  | 18 | 18 |  | 18
17 |  | 17 | 17 |  | 17
16 |  | 16 | 16 |  | 16
15 |  | 15 | 15 |  | 15
14 |  | 14 | 14 |  | 14
13 |  | 13 | 13 |  | 13
12 |  | 12 | 12 |  | 12
11 |  | 11 | 11 |  | 11
10 |  | 10 | 10 |  | 10
9 |  | 9 | 9 |  | 9
8 |  | 8 | 8 |  | 8
7 |  | 7 | 7 |  | 7
6 |  | 6 | 6 |  | 6
5 |  | 5 | 5 |  | 5
4 |  | 4 | 4 |  | 4
3 |  | 3 | 3 |  | 3
2 |  | 2 | 2 |  | 2
1 |  | 1 | 1 |  | 1

## A Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:4}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:4 |  | NA | NA | 65 | 64
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| 94 | 93 | 92 | 91 | 108 | 95 | NA | NA
CONTROLLER |  | XF:16IDA-CT{MC:3}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:3 |  | 60 | 59 | 58 | 57
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 63 | 62 | 61
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | 90 | 89 | 88 | 87 | 86 | 85 | 84
CONTROLLER |  | XF:16IDA-CT{MC:2}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:2 |  | 54 | 53 | 52 | 51
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 56 | NA | 55
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | NA | 83 | 82 | 81 | 80 | 79 | 78
CONTROLLER |  | XF:16IDA-CT{MC:1}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:1 |  | 118 | 68 | 67 | 66
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 121 | 120 | 119
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | 117 | 116 | 115 | 114 | 98 | 97 | 96 | This cables had to be separately pulled

## B Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:7}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:7 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | free | free
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | C-106 | free | free | C-18 | C-17 | C-16 | C-15
CONTROLLER |  | XF:16IDA-CT{MC:6}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:6 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-14 | C-13 | C-12 | C-11 | C-24 | C-21 | C-20 | C-19
CONTROLLER |  | XF:16IDA-CT{MC:5}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:5 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | C-23 | B-8 | B-7 | B-6
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-22 | B-10 | B-9 | B-5 | B-23 | B-22 | B-12 | B-11

## C Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:11}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:11 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | free | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | free | C-52 | C-51 | C-30 | C-29 | C-28 | C-27
CONTROLLER |  | XF:16IDA-CT{MC:10}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:10 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-33 | C-67 | C-66 | C-65 | C-26 | C-25 | C-10 | C-9
CONTROLLER |  | XF:16IDA-CT{MC:9}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
C-104 C-105
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:9 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-41 | C-40 | C-39 | C-38 | C-37 | C-36 | C-8 | C-7
CONTROLLER |  | XF:16IDA-CT{MC:8}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
C-107
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:8 |  | NA | free | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | C-45 C-46 | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-50 | C-49 | C-48 | C-32 | C-31 | free | C-35 | C-34
Special wiring