## Document information
Revision | Date | By | Changes
1 | 2014-08-08 | Oksana Ivashkevych | Original version
| 2015-03-01 | Oksana Ivashkevych
| Aug/14/2015 | Hugo Slepicka | Current version

## Component summary
Beamline Identifier | XF05ID |  | Application/GUIs
Component | Hutch | Location
LiX FOE Optics Hutch A | FE
Differential Pump | A
Beam-viewing screen #1 | A
Fixed Mask | A
Bremsstrahlung collimator 1 | A
White Beam Mirror | A
Beam-viewing screen #2 | A
White Beam Stop | A
Bremsstrahlung collimator 2 | A
Dual Multilayer Monochromator | A
Si(111) Monochromator DCM | A
Pink Beam Stop | A
Bremsstrahlung collimator 3 | A
Beam-viewing screen #3 | A
Slits | A
KB mirror system Horizontal Focusing Mirror | A
KB mirror system Vertical Focusing Mirror | A
Beam-viewing screen #4 | A
Photon Shutter | A
A
Secondry Source SOE Hutch B
Beam transport | B
Secondury Source Chamber | B
Scintilator #5 | B
Beam Position Monitor (BPM) | B
Secondary Source Chamber | B
Secondary Source Aperture(SSA) | B
Attenuator | B
Visual Beam Monitor (VBM) | B
Shutter | B
Alternative SSA | B
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | C
Harmonic Regection Mirror HRM1 | C
Transfocator CRLs | C
Harmonic Regection Mirror HRM2 | C
Divergence Defining Aperture (DDA) | C
Beam Position Monitor (BPM) | C
Guard Slits 1 in Air | C
Experimental module (user defined configuration) | C
In-air transmission measurement module
Coarse-resolution scanning | C
Fine-resolution scanning | C
Microscope | C
Guard Slits 2 in Air | C
In-vaccum GISAXS/GID module
Hexapod | C
Guard Slits 2 in VAC | C
Sample handling
Solution sample handler | C
Staubli TX40 robot | C
Detector system
WAXS detector 1 | C
WAXS detector 2 | C
WAXS beamstop | C
SAXS detector | C
SAXS beam stop | C
Note on A57: This will include multiple stages; hope to have integrated controller
Note on A58: This has its own controller

## Component naming
Beamline Identifier | 16ID
Component | Hutch | Primary subsystem | Primary instance | Primary ID | Secondary subsystem | Secondary instance | Secondary ID | Component name | Component instance | Component ID | Full component name
LiX FOE Optics Hutch A
Differential Pump | A | XF | 16IDA | XF:16IDA | BI |  | BI | DP | 1 | DP:1 | XF:16IDA-BI{DP:1}
Beam-viewing screen #1 | A | XF | 16IDA | XF:16IDA | BI |  | BI | FS | 1 | FS:1 | XF:16IDA-BI{FS:1}
Fixed Mask | A | XF | 16IDA | XF:16IDA | OP |  | OP | Msk |  | Msk | XF:16IDA-OP{Msk}
Bremsstrahlung collimator 1 | A | XF | 16IDA | XF:16IDA | OP |  | OP | BC | 1 | BC:1 | XF:16IDA-OP{BC:1}
White Beam Mirror | A | XF | 16IDA | XF:16IDA | OP |  | OP | Mir | WBM | Mir:WBM | XF:16IDA-OP{Mir:WBM}
Beam-viewing screen #2 | A | XF | 16IDA | XF:16IDA | BI |  | BI | FS | 2 | FS:2 | XF:16IDA-BI{FS:2}
White Beam Stop | A | XF | 16IDA | XF:16IDA | OP |  | OP | BS | WB | BS:WB | XF:16IDA-OP{BS:WB}
Bremsstrahlung collimator 2 | A | XF | 16IDA | XF:16IDA | OP |  | OP | BC | 2 | BC:2 | XF:16IDA-OP{BC:2}
Dual Multilayer Monochromator | A | XF | 16IDA | XF:16IDA | OP |  | OP | Mono | DMM | Mono:DMM | XF:16IDA-OP{Mono:DMM}
Si(111) Monochromator DCM | A | XF | 16IDA | XF:16IDA | OP |  | OP | Mono | DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM}
Cryocooler | A | XF | 16IDA | XF:16IDA | UT |  | UT | Cryo | 1 | Cryo:1 | XF:16IDA-UT{Cryo:1}
Pink Beam Stop | A | XF | 16IDA | XF:16IDA | OP |  | OP | BS | 1 | BS:1 | XF:16IDA-OP{BS:1}
Bremsstrahlung collimator 3 | A | XF | 16IDA | XF:16IDA | OP |  | OP | BC | 3 | BC:3 | XF:16IDA-OP{BC:3}
Beam-viewing screen #3 | A | XF | 16IDA | XF:16IDA | BI |  | BI | FS | 3 | FS:3 | XF:16IDA-BI{FS:3}
Slits | A | XF | 16IDA | XF:16IDA | OP |  | OP | Slt | 1 | Slt:1 | XF:16IDA-OP{Slt:1}
KB mirror system Horizontal Focusing Mirror | A | XF | 16IDA | XF:16IDA | OP |  | OP | Mir | KBH | Mir:KBH | XF:16IDA-OP{Mir:KBH}
KB mirror system Vertical Focusing Mirror | A | XF | 16IDA | XF:16IDA | OP |  | OP | Mir | KBV | Mir:KBV | XF:16IDA-OP{Mir:KBV}
Beam-viewing screen #4 | A | XF | 16IDA | XF:16IDA | BI |  | BI | FS | 4 | FS:4 | XF:16IDA-BI{FS:4}
Photon Shutter | A | XF | 16IDA | XF:16IDA | PPS |  | PPS | PSh |  | PSh | XF:16IDA-PPS{PSh}
Secondry Source SOE Hutch B
Beam transport | B | XF | 16IDB | XF:16IDB | OP |  | OP | BT | 1 | BT:1 | XF:16IDB-OP{BT:1}
Scintilator #5 | B | XF | 16IDB | XF:16IDB | BI |  | BI | FS | 5 | FS:5 | XF:16IDB-BI{FS:5}
Beam Position Monitor (BPM) | B | XF | 16IDB | XF:16IDB | BI |  | BI | BPM | 1 | BPM:1 | XF:16IDB-BI{BPM:1}
Secondary Source Chamber | B | XF | 16IDB | XF:16IDB | OP |  | OP | Chm | SS | Chm:SS | XF:16IDB-OP{Chm:SS}
Secondary Source Aperture(SSA) | B | XF | 16IDB | XF:16IDB | OP |  | OP | Slt | SSA | Slt:SSA | XF:16IDB-OP{Slt:SSA}
Attenuator | B | XF | 16IDB | XF:16IDB | OP |  | OP | Fltr | Atn | Fltr:Atn | XF:16IDB-OP{Fltr:Atn}
Visual Beam Monitor (VBM) | B | XF | 16IDB | XF:16IDB | BI |  | BI | FS | VBM | FS:VBM | XF:16IDB-BI{FS:VBM}
Shutter | B | XF | 16IDB | XF:16IDB | OP |  | OP | Sh |  | Sh | XF:16IDB-OP{Sh}
Alternative SSA | B | XF | 16IDB | XF:16IDB | OP |  | OP | Slt | aSSA | Slt:aSSA | XF:16IDB-OP{Slt:aSSA}
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | C | XF | 16IDC | XF:16IDC | OP |  | OP | Chm | SF | Chm:SF | XF:16IDC-OP{Chm:SF}
Harmonic Regection Mirror HRM1 | C | XF | 16IDC | XF:16IDC | OP |  | OP | Mir | HRM1 | Mir:HRM1 | XF:16IDC-OP{Mir:HRM1}
Transfocator CRLs | C | XF | 16IDC | XF:16IDC | OP |  | OP | CRL |  | CRL | XF:16IDC-OP{CRL}
Harmonic Regection Mirror HRM2 | C | XF | 16IDC | XF:16IDC | OP |  | OP | Mir | HRM2 | Mir:HRM2 | XF:16IDC-OP{Mir:HRM2}
Divergence Defining Aperture (DDA) | C | XF | 16IDC | XF:16IDC | OP |  | OP | Slt | DDA | Slt:DDA | XF:16IDC-OP{Slt:DDA}
Beam Position Monitor (BPM) | C | XF | 16IDC | XF:16IDC | BI |  | BI | BPM | 2 | BPM:2 | XF:16IDC-BI{BPM:2}
Guard Slits 1 in Air | C | XF | 16IDC | XF:16IDC | OP |  | OP | Slt | G1 | Slt:G1 | XF:16IDC-OP{Slt:G1}
Experimental module (user defined configuration) | C | XF | 16IDC | XF:16IDC | ES |  | ES | EM |  | EM | XF:16IDC-ES{EM}
In-air transmission measurement module
Coarse-resolution scanning | C | XF | 16IDC | XF:16IDC | ES | InAir | ES:InAir | Stg | ScanC | Stg:ScanC | XF:16IDC-ES:InAir{Stg:ScanC}
Fine-resolution scanning | C | XF | 16IDC | XF:16IDC | ES | InAir | ES:InAir | Stg | ScanF | Stg:ScanF | XF:16IDC-ES:InAir{Stg:ScanF}
Microscope | C | XF | 16IDC | XF:16IDC | ES | InAir | ES:InAir | Mscp | 1 | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1}
Guard Slits 2 in Air | C | XF | 16IDC | XF:16IDC | ES | InAir | ES:InAir | Slt | G2 | Slt:G2 | XF:16IDC-ES:InAir{Slt:G2}
In-vaccum GISAXS/GID module
Hexapod | C | XF | 16IDC | XF:16IDC | ES | GI | ES:GI | Stg | Hex | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex}
Guard Slits 2 in VAC | C | XF | 16IDC | XF:16IDC | ES | GI | ES:GI | Slt | G2 | Slt:G2 | XF:16IDC-ES:GI{Slt:G2}
Sample handling
Solution sample handler | C | XF | 16IDC | XF:16IDC | ES | Soln | ES:Soln | SM | Sol | SM:Sol | XF:16IDC-ES:Soln{SM:Sol}
Staubli TX40 robot | C | XF | 16IDC | XF:16IDC | ES |  | ES | SM | Staubli | SM:Staubli | XF:16IDC-ES{SM:Staubli}
Detector system
WAXS detector 1 | C | XF | 16IDC | XF:16IDC | ES |  | ES | Det | P300K1 | Det:P300K1 | XF:16IDC-ES{Det:P300K1}
WAXS detector 2 | C | XF | 16IDC | XF:16IDC | ES |  | ES | Det | P300K2 | Det:P300K2 | XF:16IDC-ES{Det:P300K2}
WAXS Detector Positioning Stage 1 | C | XF | 16IDC | XF:16IDC | ES |  | ES | Stg | WAXS1 | Stg:WAXS1 | XF:16IDC-ES{Stg:WAXS1}
WAXS Detector Positioning Stage 2 | C | XF | 16IDC | XF:16IDC | ES |  | ES | Stg | WAXS2 | Stg:WAXS2 | XF:16IDC-ES{Stg:WAXS2}
WAXS beamstop | C | XF | 16IDC | XF:16IDC | ES |  | ES | BS | WAXS | BS:WAXS | XF:16IDC-ES{BS:WAXS}
SAXS detector | C | XF | 16IDC | XF:16IDC | ES |  | ES | Det | P1M | Det:P1M | XF:16IDC-ES{Det:P1M}
SAXS Detector Positioning Stage | C | XF | 16IDC | XF:16IDC | ES |  | ES | Stg | SAXS | Stg:SAXS | XF:16IDC-ES{Stg:SAXS}
SAXS beam stop | C | XF | 16IDC | XF:16IDC | ES |  | ES | BS | SAXS | BS:SAXS | XF:16IDC-ES{BS:SAXS}

## Cameras
Component | Device name | Device Instance | Full component name | Analog cameras | Digital cameras | Position (m) | Cable ID | Camera type | Other ID | IP address | Unique ID
Hutch A
XF:16IDA-BI{FS:1} | Cam | :1 | XF:16IDA-BI{FS:1-Cam:1} |  | 1 |  | A-2 | AVT Manta G-046 1/2 |  | 10.16.1.21
XF:16IDA-BI{FS:2} | Cam | :1 | XF:16IDA-BI{FS:2-Cam:1} |  | 1 |  | A-3 | AVT Manta G-046 1/2 | 50-0503357492 (Manta_G-046C) | 10.16.1.22
XF:16IDA-BI{FS:3} | Cam | :1 | XF:16IDA-BI{FS:3-Cam:1} |  | 1 |  | A-4 | AVT Manta G-046 1/2 | 50-0503356298(Manta_G-046C) | 10.16.1.23 | 5039398
XF:16IDA-BI{FS:4} | Cam | :1 | XF:16IDA-BI{FS:4-Cam:1} |  | 1 |  | A-5 | AVT Manta G-046 1/2 | 50-0503357493 (Manta_G-046C) | 10.16.1.24
XF:16IDB-BI{FS:5} | Cam | :1 | XF:16IDB-BI{FS:5-Cam:1} |  | 1 |  | B patch panel | AVT Manta G-201-30fps 1/1.8 | 50-0503356157 (Manta_G-201C-30) | 10.16.1.25

## PLC IO summary
Component | Component ID | Hutch | Location | Temperature sensors |  | Digital IO |  |  |  | Analog IO
|  |  | (mm from source) | Pt100 | K type TC | Inputs | Type | Outputs | Type | Inputs | Type | Outputs | Type
Totals |  |  |  | 19 | 14 | 99 |  | 23 |  | 2 |  | 1
LiX FOE Optics Hutch A
Beam-viewing screen #1 | XF:16IDA-BI{FS:1} | A |  |  | 2 | 2 | Dry contact | 1 | 24V solenoid
Fixed Mask | XF:16IDA-OP{Msk} | A |  | 3
Bremsstrahlung collimator 1 | XF:16IDA-OP{BC:1} | A
White Beam Mirror | XF:16IDA-OP{Mir:WBM} | A |  | 4 | 3
Beam-viewing screen #2 | XF:16IDA-BI{FS:2} | A |  |  | 2 | 2 | Dry contact | 1 | 24V solenoid
White Beam Stop | XF:16IDA-OP{BS:WB} | A |  | 2
Bremsstrahlung collimator 2 | XF:16IDA-OP{BC:2} | A
Dual Multilayer Monochromator | XF:16IDA-OP{Mono:DMM} | A
Si(111) Monochromator DCM | XF:16IDA-OP{Mono:DCM} | #REF! |  | 8 | 4 | 4 | Dry contact |  |  |  |  | 1
Pink Beam Stop | XF:16IDA-OP{BS:1} | A |  | 2
Bremsstrahlung collimator 3 | XF:16IDA-OP{BC:3} | A |  | 0 | 0
Beam-viewing screen #3 | XF:16IDA-BI{FS:3} | A |  | 0 | 0
Slits | XF:16IDA-OP{Slt:1} | A |  | 0 | 0
KB mirror system Horizontal Focusing Mirror | XF:16IDA-OP{Mir:KBH} | A |  | 0 | 2
KB mirror system Vertical Focusing Mirror | XF:16IDA-OP{Mir:KBV} | A |  | 0 | 1
Beam-viewing screen #4 | XF:16IDA-BI{FS:4} | A |  | 0 | 0
Photon Shutter | XF:16IDA-PPS{PSh} | A
Subtotal |  |  |  | 19 | 14 | 8 |  | 2 |  | 0 |  | 1
Secondry Source SOE Hutch B
Beam transport
Scintilator #5 | FS:5 | B
Beam Position Monitor (BPM) | BPM:1 | B
Secondary Source Chamber | Chm:SS | B |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | roughing pump
Secondary Source Aperture(SSA) | Slt:SSA | B
Attenuator | Fltr:Atn | B
Visual Beam Monitor (VBM) | FS:VBM | B
Shutter | Sh | B
Alternative SSA | Slt:aSSA | B
General purpose IO |  | B
Subtotal |  |  |  | 0 | 0 | 2 |  | 1 |  | 0 |  | 0
Experimental End Station Enclosure EESE Hutch C
Secondury Focusing Chamber | Chm:SF | C |  |  |  | 2 | Dry contact | 1 | 24V solenoid |  |  |  |  | roughing pump
Secondury Focusing Chamber | Chm:SF | C
Harmonic Regection Mirror HRM1 | Mir:HRM1 | C
Transfocator CRLs | CRL | C |  |  |  | 27 |  | 9
Harmonic Regection Mirror HRM2 | Mir:HRM2 | C | [7 empty] | 2 |  |  |  | from LVDT sensor, additional module integration
Divergence Defining Aperture (DDA) | Slt:DDA | C
Beam Position Monitor (BPM) | BPM:2 | C
Guard Slits 1 in Air | Slt:G1 | C
Experimental module (user defined configuration) | EM | C |  |  |  | 2 | Dry contact; signal about type of ES | [6 empty] | roughing pump
In-air transmission measurement module
Coarse-resolution scanning | Stg:ScanC | C
Fine-resolution scanning | Stg:ScanF | C
Microscope | Mscp:1 | C
Guard Slits 2 in Air | Slt:G2 | C
In-vaccum GISAXS/GID module
Hexapod | Stg:Hex | C
Guard Slits 2 in VAC | Slt:G2 | C
Sample handling
Solution sample handler | SM:Sol | C
Staubli TX40 robot | SM:Staubli | C
Detector system
WAXS detector 1 | Det:P300K1 | C
WAXS detector 2 | Det:P300K2 | C
WAXS beamstop | BS:WAXS | C
SAXS detector | Det:P1M | C
SAXS beam stop | BS:SAXS | C
General purpose IO |  | C
Subtotal |  |  |  | 0 | 0 | 31 |  | 10 |  | 2 |  | 0
Vacuum system
XF:16IDA-VA{FE-GV:1} |  | A | 0 |  |  | 0 |  | 0 | [5 empty] | Handled by front end
XF:16IDA-VA{DP:1-IP:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{DP:1-CCG:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{DP:1-TCG:1} |  | A | 26.3 |  |  | 1
XF:16IDA-VA{DP:1-IP:2} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{DP:1-CCG:2} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{DP:1-TCG:2} |  | A | 28.2 |  |  | 1
XF:16IDA-VA{DP:1-GV:1} |  | A | 28.2 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mir:WBM-IP:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-CCG:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-TCG:1} |  | A | 29.1 |  |  | 1
XF:16IDA-VA{Mir:WBM-GV:1} |  | A | 29.1 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mono:DMM-IP:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:DMM-CCG:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:DMM-TCG:1} |  | A | 30.7 |  |  | 1
XF:16IDA-VA{Mono:DMM-IP:2} |  | A | 31.5 |  |  | 1
XF:16IDA-VA{Mono:DMM-CCG:1} |  | A | 31.5 |  |  | 1
XF:16IDA-VA{Cryo:1-TCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:DMM-GV:1} |  | A | 32.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mono:DCM-IP:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:DCM-CCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:DCM-TCG:1} |  | A | 32.7 |  |  | 1
XF:16IDA-VA{Mono:DCM-GV:1} |  | A | 32.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{FS:3-IP:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{FS:3-CCG:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{FS:3-TCG:1} |  | A | 36.3 |  |  | 1
XF:16IDA-VA{FS:3-GV:1} |  | A | 36.3 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{Mir:KB-IP:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-CCG:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-TCG:1} |  | A | 37.7 |  |  | 1
XF:16IDA-VA{Mir:KB-GV:1} |  | A | 37.7 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDA-VA{BT:1-IP:1} |  | A | 39 |  |  | 1
XF:16IDA-VA{BT:1-CCG:1} |  | A | 39 |  |  | 1
XF:16IDA-VA{BT:1-TCG:1} |  | A | 39 |  |  | 1
XF:16IDB-VA{BT:1-IP:2} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{BT:1-CCG:2} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{BT:1-TCG:1} |  | B | 50.1 |  |  | 1
XF:16IDB-VA{BT:1-GV:1} |  | B | 50.1 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDB-VA{Chm:SS-TCG:1} |  | B | 51 |  |  | 1
XF:16IDB-VA{Chm:SS-TCG:2} |  | B | 51 |  |  | 1
XF:16IDC-VA{Chm:SF-GV:1} |  | C | 51 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{Chm:SF-TCG:1} |  | C | 58.875 |  |  | 1
XF:16IDC-VA{Chm:SF-GV:1} |  | C | 58.9 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{EM-TCG:1} |  | C | 59.5 |  |  | 1
XF:16IDC-VA{EM-GV:1} |  | C | 59.5 |  |  | 2 |  | 1 | 12 V solenoid
XF:16IDC-VA{det:WAXS-TCG:1} |  | C | 60 |  |  | 1
Equipment identification signal |  | C | 60 |  |  | 3
Subtotal |  |  |  | 0 | 0 | 58 |  | 10 |  | 0 |  | 0
Cooling water
Beam-viewing screen #1 | FS:1 | A | Roof | [6 empty] | 1 | 4-20 mA
Fixed Mask | Msk | A | Roof | [6 empty] | 1 | 4-20 mA
White Beam Stop | BS:WB | A | Roof | [6 empty] | 1 | 4-20 mA
Beam-viewing screen #2 | FS:2 | A | Roof | [6 empty] | 1 | 4-20 mA
Pink Beam Stop | BS:1 | A | Roof | [6 empty] | 1 | 4-20 mA
Subtotal |  |  |  | 0 | 0 | 0 |  | 0 |  | 5 |  | 0
Total |  |  |  | 19 | 14 | 99 |  | 23 |  | 2 |  | 1
Other (So far irrelevant) Info below is from SRX
Synchrotron beam not inhibited by PPS | [5 empty] | 1 | Scott Buda to define
Synchrotron beam not inhibited by EPS | [7 empty] | 1 | Scott Buda to define
Cooling water supply isolation valve |  | A | Upstream |  |  | 2 |  | 2
Cooling water return isolation valve |  | A | Upstream |  |  | 2 |  | 2
Leak detection system |  | A | ? |  |  | 2?
Cooling water supply temperature |  | A | Upstream | 1
Cooling water return temperature |  | A | Upstream | 1
Cryocooler |  | A | Mezzanine |  |  | 1
Cryocooler disable |  | A | Mezzanine |  |  |  |  | 1
A Hutch O2 sensor |  | A | ? |  |  | 1
D Hutch O2 sensor |  | D | ? |  |  | 1
N2 system disable |  | A | ? |  |  |  |  | 1
Motion controller disable |  | A | RG:A2 |  |  |  |  | 4
Motion controller disable |  | D | RG:D1? |  |  |  |  | 4
D Hutch endstation motion interlocks |  | D | 65000 |  |  | 20 | Dry contact | 10 | 24 VDC solenoid
A Hutch PPS |  | A |  |  |  | 3 |  | 3
B Hutch PPS |  | B |  |  |  | 3 |  | 2
D Hutch PPS |  | D |  |  |  | 3 |  | 2
Subtotal |  |  |  | 2 | 0 | 37 |  | 32 |  | 0 |  | 0
Total |  |  |  | #REF! | #REF! | #REF! |  | #REF! |  | #REF! |  | #REF!
Note on A123: Ivashkevych, Oksana: this is to identify which component is installed, up to 2^3; some logig will be involved, you have to first disconnect A before connecting B

## PLC IO Detail
[6 empty] | Fault condition | Hi Alarm [°C] | Trip Point [°C]
Component | Measurement ( Description) | Signal type | Device | Component name | Signal name |  |  |  | Signal Name | Output | Action | Output | Action
A HUTCH
Beam-viewing screen #1 | Upper Inboard | analog input | K-Type | XF:16IDA-BI{FS:1} | T:1-I |  | 90 | 100 | XF:16IDA-BI{FS:1}T:1-I | FE Shutter | Close
Beam-viewing screen #1 | Lower Outboard | analog input | K-Type | XF:16IDA-BI{FS:1} | T:2-I |  | 90 | 100 | XF:16IDA-BI{FS:1}T:2-I | FE Shutter | Close
Beam-viewing screen #1 | Staus In | digital input | dry contact | XF:16IDA-BI{FS:1} | Pos:In-Sts |  |  |  | XF:16IDA-BI{FS:1}Pos:In-Sts
Beam-viewing screen #1 | Status Out | digital input | dry contact | XF:16IDA-BI{FS:1} | Pos:Out-Sts |  |  |  | XF:16IDA-BI{FS:1}Pos:Out-Sts
Beam-viewing screen #1 | Command | digital output | 24V solenoid | XF:16IDA-BI{FS:1} | Cmd |  |  |  | XF:16IDA-BI{FS:1}Cmd
Fixed Mask | Top Upstream | analog input | RTD | XF:16IDA-OP{Msk} | T:1-I |  | 90 | 100 | XF:16IDA-OP{Msk}T:1-I | FE Shutter | Close
Fixed Mask | Top Downstream | analog input | RTD | XF:16IDA-OP{Msk} | T:2-I |  | 90 | 100 | XF:16IDA-OP{Msk}T:2-I | FE Shutter | Close
Fixed Mask | Bottom Downstream | analog input | RTD | XF:16IDA-OP{Msk} | T:3-I |  | 90 | 100 | XF:16IDA-OP{Msk}T:3-I | FE Shutter | Close
White Beam Mirror |  | analog input | flow meter | XF:16IDA-OP{Mir:WBM} | F-I |  | 0.4 GPM | 0.3 GPM | XF:16IDA-OP{Mir:WBM}F-I
White Beam Mirror | Mask Left | analog input | RTD | XF:16IDA-OP{Mir:WBM} | T:Msk1-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM}T:Msk1-I | FE Shutter | Close
White Beam Mirror | Mask Right | analog input | RTD | XF:16IDA-OP{Mir:WBM} | T:Msk2-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM}T:Msk2-I | FE Shutter | Close
White Beam Mirror | Mirr downstream | analog input | RTD | XF:16IDA-OP{Mir:WBM} | T:D-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM}T:D-I | FE Shutter | Close
White Beam Mirror | Mirr leteral | analog input | RTD | XF:16IDA-OP{Mir:WBM} | T:L-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM}T:L-I | FE Shutter | Close
White Beam Stop | Bottom Upstream | analog input | RTD | XF:16IDA-OP{BS:WB} | T:1-I |  | 90 | 100 | XF:16IDA-OP{BS:WB}T:1-I | FE Shutter | Close
White Beam Stop | Bottom Downstream | analog input | RTD | XF:16IDA-OP{BS:WB} | T:2-I |  | 90 | 100 | XF:16IDA-OP{BS:WB}T:2-I | FE Shutter | Close
White Beam Mirror | On the motor | analog input | K-Type | XF:16IDA-OP{Mir:WBM-Ax:X} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM-Ax:X}T-I |  | Kill PWR mc01
White Beam Mirror | On the motor | analog input | K-Type | XF:16IDA-OP{Mir:WBM-Ax:Y} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM-Ax:Y}T-I |  | Kill PWR mc01
White Beam Mirror | On the motor | analog input | K-Type | XF:16IDA-OP{Mir:WBM-Ax:P} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:WBM-Ax:P}T-I |  | Kill PWR mc01
Beam-viewing screen #2 | Inboard Upper | analog input | K-Type | XF:16IDA-BI{FS:2} | T:1-I |  | 90 | 100 | XF:16IDA-BI{FS:2}T:1-I | FE Shutter | Close
Beam-viewing screen #2 | Outboard Lower | analog input | K-Type | XF:16IDA-BI{FS:2} | T:2-I |  | 90 | 100 | XF:16IDA-BI{FS:2}T:2-I | FE Shutter | Close
Beam-viewing screen #2 | Staus In | digital input | dry contact | XF:16IDA-BI{FS:2} | Pos:In-Sts |  |  |  | XF:16IDA-BI{FS:2}Pos:In-Sts
Beam-viewing screen #2 | Status Out | digital input | dry contact | XF:16IDA-BI{FS:2} | Pos:Out-Sts |  |  |  | XF:16IDA-BI{FS:2}Pos:Out-Sts
Beam-viewing screen #2 | Command | digital output | 24V solenoid | XF:16IDA-BI{FS:2} | Cmd |  |  |  | XF:16IDA-BI{FS:2}Cmd
Si(111) Monochromator DCM | On FE.Y motor | analog input | K-Type | XF:16IDA-OP{Mono:DCM-Ax:Of2} | T-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM-Ax:Of2}T-I |  | Kill PWR mc02
Si(111) Monochromator DCM | ON FE.P motor | analog input | K-Type | XF:16IDA-OP{Mono:DCM-Ax:P2} | T-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM-Ax:P2}T-I |  | Kill PWR mc02
Si(111) Monochromator DCM | On FE.R motor | analog input | K-Type | XF:16IDA-OP{Mono:DCM-Ax:R2} | T-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM-Ax:R2}T-I |  | Kill PWR mc02
Si(111) Monochromator DCM | On CC.P motor | analog input | K-Type | XF:16IDA-OP{Mono:DCM-Ax:CCM_PF} | T-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM-Ax:CCM_PF}T-I |  | Kill PWR mc02
Si(111) Monochromator DCM | Loc A Second crystal bulk | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:A1-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:A1-I
Si(111) Monochromator DCM | Loc A Second crystal leteral absorber | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:A2-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:A2-I
Si(111) Monochromator DCM | Loc A Second Crystal Pitch Flexture | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:A3-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:A3-I
Si(111) Monochromator DCM | Loc B Detuning Compton Shield | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:B1-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:B1-I
Si(111) Monochromator DCM | Loc B Front Mask DCM | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:B2-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:B2-I
Si(111) Monochromator DCM | Loc B Front mask CC | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:B3-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:B3-I
Si(111) Monochromator DCM | Loc B First Crystal bulk | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:B4-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:B4-I
Si(111) Monochromator DCM | Loc B Warm Plate | analog input | RTD | XF:16IDA-OP{Mono:DCM} | T:B5-I |  | 80 | 90 | XF:16IDA-OP{Mono:DCM}T:B5-I
Si(111) Monochromator DCM | CW Bragg Safety Limit | digital input | limit switch | XF:16IDA-OP{Mono:DCM-Ax:X} | Sw:OvrP-Sts |  |  |  | XF:16IDA-OP{Mono:DCM-Ax:X}Sw:OvrP-Sts |  | Kill PWR mc02
Si(111) Monochromator DCM | CCW Bragg Safety Limit | digital input | limit switch | XF:16IDA-OP{Mono:DCM-Ax:X} | Sw:OvrN-Sts |  |  |  | XF:16IDA-OP{Mono:DCM-Ax:X}Sw:OvrN-Sts |  | Kill PWR mc02
Si(111) Monochromator DCM | DCM Positin | digital input | limit switch | XF:16IDA-OP{Mono:DCM} | DCM_Pos-Sts |  |  |  | XF:16IDA-OP{Mono:DCM}DCM_Pos-Sts
Si(111) Monochromator DCM | CCM Position | digital input | limit switch | XF:16IDA-OP{Mono:DCM} | CCM_Pos-Sts |  |  |  | XF:16IDA-OP{Mono:DCM}CCM_Pos-Sts
LN2 valve main | Staus In | digital input | dry contact | XF:16IDA-OP{LN2-Vlv} | Pos:In-Sts |  |  |  | XF:16IDA-OP{LN2-Vlv}Pos:In-Sts
LN2 valve main | Status Out | digital input | dry contact | XF:16IDA-OP{LN2-Vlv} | Pos:Out-Sts |  |  |  | XF:16IDA-OP{LN2-Vlv}Pos:Out-Sts
LN2 valve main | Command | digital output | 24V solenoid | XF:16IDA-OP{LN2-Vlv} | Cmd |  |  |  | XF:16IDA-OP{LN2-Vlv}Cmd
LN2 valve local | Staus In | digital input | dry contact | XF:16IDA-UT{Cryo:1} | Pos:In-Sts |  |  |  | XF:16IDA-UT{Cryo:1}Pos:In-Sts
LN2 valve local | Status Out | digital input | dry contact | XF:16IDA-UT{Cryo:1} | Pos:Out-Sts |  |  |  | XF:16IDA-UT{Cryo:1}Pos:Out-Sts
LN2 valve local | Command | digital output | 24V solenoid | XF:16IDA-UT{Cryo:1} | Cmd |  |  |  | XF:16IDA-UT{Cryo:1}Cmd
Pink Beam Stop | Bottom Upstream | analog input | RTD | XF:16IDA-OP{BS:1} | T:1-I |  | 90 | 100 | XF:16IDA-OP{BS:1}T:1-I | FE Shutter | Close
Pink Beam Stop | Bottom Downstream | analog input | RTD | XF:16IDA-OP{BS:1} | T:2-I |  | 90 | 100 | XF:16IDA-OP{BS:1}T:2-I | FE Shutter | Close
KB mirror system Horizontal Focusing Mirror | On HFMX-upstream mtr | analog input | K-Type | XF:16IDA-OP{Mir:KBH-Ax:XU} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:KBH-Ax:XU}T-I |  | Kill PWR mc03
KB mirror system Horizontal Focusing Mirror | On HFMX-downstream mtr | analog input | K-Type | XF:16IDA-OP{Mir:KBH-Ax:XD} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:KBH-Ax:XD}T-I |  | Kill PWR mc03
KB mirror system Vertical Focusing Mirror | On VFMX horizontal translation | analog input | K-Type | XF:16IDA-OP{Mir:KBV-Ax:X} | T-I |  | 80 | 90 | XF:16IDA-OP{Mir:KBV-Ax:X}T-I |  | Kill PWR mc03
Cryocooler
|  | analog input |  | XF:16IDA-UT{Cryo:1} | Sts:OK-Sts |  | Fault | Fault | XF:16IDA-UT{Cryo:1}Sts:OK-Sts
|  | analog input |  | XF:16IDA-UT{Cryo:1} | Rdy-Sts |  | Off | Off | XF:16IDA-UT{Cryo:1}Rdy-Sts
|  | digital output |  | XF:16IDA-UT{Cryo:1} | Cmd:Run-Cmd |  |  |  | XF:16IDA-UT{Cryo:1}Cmd:Run-Cmd
C HUTCH
Harmonic Regection Mirror HRM2 | Upstream Bend; with LVDT | analog input | LVDT | XF:16IDC-OP{Mir:HRM2-Ax:BU} | LVDT |  |  |  | XF:16IDC-OP{Mir:HRM2-Ax:BU}LVDT
Harmonic Regection Mirror HRM2 | Downstream Bend; With LVDT | analog input | LVDT | XF:16IDC-OP{Mir:HRM2-Ax:BD} | LVDT |  |  |  | XF:16IDC-OP{Mir:HRM2-Ax:BD}LVDT
Transfocator | CRL system | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:Sys} | Enbl |  |  |  | XF:16IDC-OP{CRL:Sys}Enbl
Transfocator | CRL 1 | TTL digital input | 5V input | XF:16IDC-OP{CRL:1} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:1}Pos:In-Sts
Transfocator | CRL 1 | TTL digital input | 5V input | XF:16IDC-OP{CRL:1} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:1}Pos:Out-Sts
Transfocator | CRL 1 | TTL digital input | 5V input | XF:16IDC-OP{CRL:1} | Enbl |  |  |  | XF:16IDC-OP{CRL:1}Enbl
Transfocator | CRL 1 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:1} | Cmd |  |  |  | XF:16IDC-OP{CRL:1}Cmd
Transfocator | CRL 2 | TTL digital input | 5V input | XF:16IDC-OP{CRL:2} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:2}Pos:In-Sts
Transfocator | CRL 2 | TTL digital input | 5V input | XF:16IDC-OP{CRL:2} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:2}Pos:Out-Sts
Transfocator | CRL 2 | TTL digital input | 5V input | XF:16IDC-OP{CRL:2} | Enbl |  |  |  | XF:16IDC-OP{CRL:2}Enbl
Transfocator | CRL 2 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:2} | Cmd |  |  |  | XF:16IDC-OP{CRL:2}Cmd
Transfocator | CRL 3 | TTL digital input | 5V input | XF:16IDC-OP{CRL:3} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:3}Pos:In-Sts
Transfocator | CRL 3 | TTL digital input | 5V input | XF:16IDC-OP{CRL:3} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:3}Pos:Out-Sts
Transfocator | CRL 3 | TTL digital input | 5V input | XF:16IDC-OP{CRL:3} | Enbl |  |  |  | XF:16IDC-OP{CRL:3}Enbl
Transfocator | CRL 3 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:3} | Cmd |  |  |  | XF:16IDC-OP{CRL:3}Cmd
Transfocator | CRL 4 | TTL digital input | 5V input | XF:16IDC-OP{CRL:4} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:4}Pos:In-Sts
Transfocator | CRL 4 | TTL digital input | 5V input | XF:16IDC-OP{CRL:4} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:4}Pos:Out-Sts
Transfocator | CRL 4 | TTL digital input | 5V input | XF:16IDC-OP{CRL:4} | Enbl |  |  |  | XF:16IDC-OP{CRL:4}Enbl
Transfocator | CRL 4 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:4} | Cmd |  |  |  | XF:16IDC-OP{CRL:4}Cmd
Transfocator | CRL 5 | TTL digital input | 5V input | XF:16IDC-OP{CRL:5} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:5}Pos:In-Sts
Transfocator | CRL 5 | TTL digital input | 5V input | XF:16IDC-OP{CRL:5} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:5}Pos:Out-Sts
Transfocator | CRL 5 | TTL digital input | 5V input | XF:16IDC-OP{CRL:5} | Enbl |  |  |  | XF:16IDC-OP{CRL:5}Enbl
Transfocator | CRL 5 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:5} | Cmd |  |  |  | XF:16IDC-OP{CRL:5}Cmd
Transfocator | CRL 6 | TTL digital input | 5V input | XF:16IDC-OP{CRL:6} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:6}Pos:In-Sts
Transfocator | CRL 6 | TTL digital input | 5V input | XF:16IDC-OP{CRL:6} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:6}Pos:Out-Sts
Transfocator | CRL 6 | TTL digital input | 5V input | XF:16IDC-OP{CRL:6} | Enbl |  |  |  | XF:16IDC-OP{CRL:6}Enbl
Transfocator | CRL 6 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:6} | Cmd |  |  |  | XF:16IDC-OP{CRL:6}Cmd
Transfocator | CRL 7 | TTL digital input | 5V input | XF:16IDC-OP{CRL:7} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:7}Pos:In-Sts
Transfocator | CRL 7 | TTL digital input | 5V input | XF:16IDC-OP{CRL:7} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:7}Pos:Out-Sts
Transfocator | CRL 7 | TTL digital input | 5V input | XF:16IDC-OP{CRL:7} | Enbl |  |  |  | XF:16IDC-OP{CRL:7}Enbl
Transfocator | CRL 7 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:7} | Cmd |  |  |  | XF:16IDC-OP{CRL:7}Cmd
Transfocator | CRL 8 | TTL digital input | 5V input | XF:16IDC-OP{CRL:8} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:8}Pos:In-Sts
Transfocator | CRL 8 | TTL digital input | 5V input | XF:16IDC-OP{CRL:8} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:8}Pos:Out-Sts
Transfocator | CRL 8 | TTL digital input | 5V input | XF:16IDC-OP{CRL:8} | Enbl |  |  |  | XF:16IDC-OP{CRL:8}Enbl
Transfocator | CRL 8 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:8} | Cmd |  |  |  | XF:16IDC-OP{CRL:8}Cmd
Transfocator | CRL 9 | TTL digital input | 5V input | XF:16IDC-OP{CRL:9} | Pos:In-Sts |  |  |  | XF:16IDC-OP{CRL:9}Pos:In-Sts
Transfocator | CRL 9 | TTL digital input | 5V input | XF:16IDC-OP{CRL:9} | Pos:Out-Sts |  |  |  | XF:16IDC-OP{CRL:9}Pos:Out-Sts
Transfocator | CRL 9 | TTL digital input | 5V input | XF:16IDC-OP{CRL:9} | Enbl |  |  |  | XF:16IDC-OP{CRL:9}Enbl
Transfocator | CRL 9 | digital relay output | 24 V drycontact output | XF:16IDC-OP{CRL:9} | Cmd |  |  |  | XF:16IDC-OP{CRL:9}Cmd
Venting Valves
Bypass Valve 1 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-BV:1} | Sts:Open |  |  |  | XF:16IDC-VA{ES-BV:1}Sts:Open
Bypass Valve 1 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-BV:1} | Sts:Close |  |  |  | XF:16IDC-VA{ES-BV:1}Sts:Close
Evacuate Valve 1 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:1} | Sts:Open |  |  |  | XF:16IDC-VA{ES-EV:1}Sts:Open
Evacuate Valve 1 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:1} | Sts:Close |  |  |  | XF:16IDC-VA{ES-EV:1}Sts:Close
Evacuate Valve 2 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:2} | Sts:Open |  |  |  | XF:16IDC-VA{ES-EV:2}Sts:Open
Evacuate Valve 2 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:2} | Sts:Close |  |  |  | XF:16IDC-VA{ES-EV:2}Sts:Close
Evacuate Valve 3 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:3} | Sts:Open |  |  |  | XF:16IDC-VA{ES-EV:3}Sts:Open
Evacuate Valve 3 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:3} | Sts:Close |  |  |  | XF:16IDC-VA{ES-EV:3}Sts:Close
Evacuate Valve 4 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:4} | Sts:Open |  |  |  | XF:16IDC-VA{ES-EV:4}Sts:Open
Evacuate Valve 4 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-EV:4} | Sts:Close |  |  |  | XF:16IDC-VA{ES-EV:4}Sts:Close
Venting Valve 1 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:1} | Sts:Open |  |  |  | XF:16IDC-VA{ES-VV:1}Sts:Open
Venting Valve 1 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:1} | Sts:Close |  |  |  | XF:16IDC-VA{ES-VV:1}Sts:Close
Venting Valve 2 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:2} | Sts:Open |  |  |  | XF:16IDC-VA{ES-VV:2}Sts:Open
Venting Valve 2 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:2} | Sts:Close |  |  |  | XF:16IDC-VA{ES-VV:3}Sts:Close
Venting Valve 3 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:3} | Sts:Open |  |  |  | XF:16IDC-VA{ES-VV:3}Sts:Open
Venting Valve 3 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:3} | Sts:Close |  |  |  | XF:16IDC-VA{ES-VV:3}Sts:Close
Venting Valve 4 | valve open | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:4} | Sts:Open |  |  |  | XF:16IDC-VA{ES-VV:4}Sts:Open
Venting Valve 4 | valve closed | digital input | 24V solenoid | XF:16IDC-VA{ES-VV:4} | Sts:Close |  |  |  | XF:16IDC-VA{ES-VV:4}Sts:Close

## PLC EPS IO Logic
Input | Signal type | Comparison | Setting | Set Point Value, mbar | AND input | AND comparison | AND value | Output | Action | Output | Action
XF:16IDA-OP{Mono:DCM}DCM_Pos-Sts | digital input | != | active |  | XF:16IDA-OP{Mono:DCM}CCM_Pos-Sts | != | active | FE Shutter | Close
XF:16IDA-OP{Mono:DCM}CCM_Pos-Sts | digital input | = | active |  | XF:16IDA-OP{BS:1}T:1-I | = | active | FE Shutter | Close | raise error somehow
Current in the ring |  | > | 25 mA |  |  |  |  | XF:16IDA-BI{FS:1}Cmd | Out
Power to the Cryocoller |  | = | ON |  |  |  |  | XF:16IDA-OP{LN2-Vlv}Cmd | Open | XF:16IDA-UT{Cryo:1}Cmd | Open
Power to the Cryocoller |  | = | OFF |  |  |  |  | Send email to list

## Vacuum components
16ID
Vac. Sect. | Hutch | Location (meters) | Primary subsystem | Primary instance | Primary ID | Secondary subsystem | Branch | Secondary ID | Primary device name | Device name | Device instance | Device ID | Full component name | Associated component | Controller | Channel | Cable number | Cable number
LINK
1 | A |  | XF | 16IDA | XF:16IDA | VA |  | VA | FE | GV | 1 | GV:1 | XF:16IDA-VA{FE-GV:1}
| A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | IP | 1 | IP:1 | XF:16IDA-VA{DP:1-IP:1} |  | XF:16IDA-CT{IPC:1} | 1 | 74 | 74-IP
| A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | CCG | 1 | CCG:1 | XF:16IDA-VA{DP:1-CCG:1} |  | XF:16IDA-CT{VGC:1} | 1 | 47 | 47-CCG SIG
| A | 26.3 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | TCG | 1 | TCG:1 | XF:16IDA-VA{DP:1-TCG:1} |  | XF:16IDA-CT{VGC:1} | 5 | 99 | 99-PIRANI
| A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | IP | 2 | IP:2 | XF:16IDA-VA{DP:1-IP:2} |  | XF:16IDA-CT{IPC:1} | 2 | 69 | 69-IP
| A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | CCG | 2 | CCG:2 | XF:16IDA-VA{DP:1-CCG:2} |  | XF:16IDA-CT{VGC:1} | 3 | 42 | 42-CCG SIG
| A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | TCG | 2 | TCG:2 | XF:16IDA-VA{DP:1-TCG:2} |  | XF:16IDA-CT{VGC:1} | 6 | 100 | 100-PIRANI
2 | A | 28.2 | XF | 16IDA | XF:16IDA | VA |  | VA | DP:1 | GV | 1 | GV:1 | XF:16IDA-VA{DP:1-GV:1}
| A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | IP | 1 | IP:1 | XF:16IDA-VA{Mir:WBM-IP:1} |  | XF:16IDA-CT{IPC:2} | 1 | 77 | 77-IP
| A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | CCG | 1 | CCG:1 | XF:16IDA-VA{Mir:WBM-CCG:1} |  | XF:16IDA-CT{VGC:2} | 1 | 50 | 50-CCG SIG
| A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | TCG | 1 | TCG:1 | XF:16IDA-VA{Mir:WBM-TCG:1} |  | XF:16IDA-CT{VGC:2} | 5 | 107 | 107-PIRANI
3 | A | 29.1 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:WBM | GV | 1 | GV:1 | XF:16IDA-VA{Mir:WBM-GV:1}
| A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | IP | 1 | IP:1 | XF:16IDA-VA{Mono:DMM-IP:1} |  | XF:16IDA-CT{IPC:2} | 2 | 72 | 72-IP
| A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:DMM-CCG:1} |  | XF:16IDA-CT{VGC:2} | 3 | 45 | 45-CCG SIG
| A | 30.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | TCG | 1 | TCG:1 | XF:16IDA-VA{Mono:DMM-TCG:1} |  | XF:16IDA-CT{VGC:2} | 6 | 103 | 103-PIRANI
| A | 31.5 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | IP | 2 | IP:2 | XF:16IDA-VA{Mono:DMM-IP:2} |  | XF:16IDA-CT{IPC:3} | 1 | 73 | IP-73
| A | 31.5 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:DMM-CCG:1} |  | XF:16IDA-CT{VGC:3} | 1 | 46 | no place
XXXX | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Cryo:1 | TCG | 1 | TCG:1 | XF:16IDA-VA{Cryo:1-TCG:1} |  | XF:16IDA-CT{VGC:3} | 5 | 104 | 104-PIRANI
4 | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DMM | GV | 1 | GV:1 | XF:16IDA-VA{Mono:DMM-GV:1}
| A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DCM | IP | 1 | IP:1 | XF:16IDA-VA{Mono:DCM-IP:1} |  | XF:16IDA-CT{IPC:3} | 2 | 70 | 70-IP
| A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DCM | CCG | 1 | CCG:1 | XF:16IDA-VA{Mono:DCM-CCG:1} |  | XF:16IDA-CT{VGC:3} | 3 | 43 | 43-CCG SIG
| A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DCM | TCG | 1 | TCG:1 | XF:16IDA-VA{Mono:DCM-TCG:1} |  | XF:16IDA-CT{VGC:3} | 6 | 101 | 101-PIRANI
5 | A | 32.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mono:DCM | GV | 1 | GV:1 | XF:16IDA-VA{Mono:DCM-GV:1}
| A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | FS:3 | IP | 1 | IP:1 | XF:16IDA-VA{FS:3-IP:1} |  | XF:16IDA-CT{IPC:4} | 1 | 75 | 75-IP
| A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | FS:3 | CCG | 1 | CCG:1 | XF:16IDA-VA{FS:3-CCG:1} |  | XF:16IDA-CT{VGC:4} | 1 | 48 | 48-CCG SIG
| A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | FS:3 | TCG | 1 | TCG:1 | XF:16IDA-VA{FS:3-TCG:1} |  | XF:16IDA-CT{VGC:4} | 5 | 105 | 105-PIRANI
6 | A | 36.3 | XF | 16IDA | XF:16IDA | VA |  | VA | FS:3 | GV | 1 | GV:1 | XF:16IDA-VA{FS:3-GV:1}
| A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | IP | 1 | IP:1 | XF:16IDA-VA{Mir:KB-IP:1} |  | XF:16IDA-CT{IPC:4} | 2 | 71 | 71-IP
| A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | CCG | 1 | CCG:1 | XF:16IDA-VA{Mir:KB-CCG:1} |  | XF:16IDA-CT{VGC:3} | 1 | 44 | 44-CCG SIG
| A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | TCG | 1 | TCG:1 | XF:16IDA-VA{Mir:KB-TCG:1} |  | XF:16IDA-CT{VGC:4} | 6 | 102 | 102-PIRANI
| A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | Trip | HV | Trip:HV | XF:16IDA-VA{Mir:KB-Trip:HV} |  | XF:16IDA-CT{VGC:4}
7 | A | 37.7 | XF | 16IDA | XF:16IDA | VA |  | VA | Mir:KB | GV | 1 | GV:1 | XF:16IDA-VA{Mir:KB-GV:1}
| A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | BT:1 | IP | 1 | IP:1 | XF:16IDA-VA{BT:1-IP:1} |  | XF:16IDA-CT{IPC:5} | 1 | 76 | 76-IP
| A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | BT:1 | CCG | 1 | CCG:1 | XF:16IDA-VA{BT:1-CCG:1} |  | XF:16IDA-CT{VGC:4} | 2 | 49 | 49-CCG SIG
| A | 39 | XF | 16IDA | XF:16IDA | VA |  | VA | BT:1 | TCG | 1 | TCG:1 | XF:16IDA-VA{BT:1-TCG:1} |  | XF:16IDA-CT{VGC:4} | 6 | 106
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 1 | VGC:1 | XF:16IDA-CT{VGC:1}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 2 | VGC:2 | XF:16IDA-CT{VGC:2}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 3 | VGC:3 | XF:16IDA-CT{VGC:3}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | VGC | 4 | VGC:4 | XF:16IDA-CT{VGC:4}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 1 | IPC:1 | XF:16IDA-CT{IPC:1}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 2 | IPC:2 | XF:16IDA-CT{IPC:2}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 3 | IPC:3 | XF:16IDA-CT{IPC:3}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 4 | IPC:4 | XF:16IDA-CT{IPC:4}
| A | RG:A3 | XF | 16IDA | XF:16IDA | CT |  | CT |  | IPC | 5 | IPC:5 | XF:16IDA-CT{IPC:5}
| B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | BT:1 | IP | 2 | IP:2 | XF:16IDB-VA{BT:1-IP:2} |  | XF:16IDA-CT{IPC:5} | 2 | 3
| B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | BT:1 | CCG | 2 | CCG:2 | XF:16IDB-VA{BT:1-CCG:2} |  | XF:16IDB-CT{VGC:1} | 1 | 2
| B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | BT:1 | TCG | 1 | TCG:1 | XF:16IDB-VA{BT:1-TCG:1} |  | XF:16IDB-CT{VGC:1} | 5 | 24
END station Vacuum an Venting Section
8 | B | 50.1 | XF | 16IDB | XF:16IDB | VA |  | VA | BT:1 | GV | 1 | GV:1 | XF:16IDB-VA{BT:1-GV:1} | ES_GV0
| B | 51 | XF | 16IDB | XF:16IDB | VA |  | VA | Chm:SS | TCG | 1 | TCG:1 | XF:16IDB-VA{Chm:SS-TCG:1} |  | XF:16IDB-CT{VGC:1} | 6 | 25
| B | 51 | XF | 16IDB | XF:16IDB | VA |  | VA | Chm:SS | TCG | 2 | TCG:2 | XF:16IDB-VA{Chm:SS-TCG:2} |  | XF:16IDB-CT{VGC:2} | 1 | 38
9 | C | 51 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | GV | 1 | GV:1 | XF:16IDC-VA{Chm:SF-GV:1} | ES_GV1
| C | 58.875 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | TCG | 1 | TCG:1 | XF:16IDC-VA{Chm:SF-TCG:1} |  | XF:16IDB-CT{VGC:2} | 2 | 68
10 | C | 58.9 | XF | 16IDC | XF:16IDC | VA |  | VA | Chm:SF | GV | 1 | GV:1 | XF:16IDC-VA{Chm:SF-GV:1} | ES_GV2
| C | 59.5 | XF | 16IDC | XF:16IDC | VA |  | VA | EM | TCG | 1 | TCG:1 | XF:16IDC-VA{EM-TCG:1} |  | XF:16IDB-CT{VGC:2} | 3 | 69
11 | C | 59.5 | XF | 16IDC | XF:16IDC | VA |  | VA | EM | GV | 1 | GV:1 | XF:16IDC-VA{EM-GV:1} | ES_GV3
| C | 60 | XF | 16IDC | XF:16IDC | VA |  | VA | det:WAXS | TCG | 1 | TCG:1 | XF:16IDC-VA{det:WAXS-TCG:1} |  | XF:16IDB-CT{VGC:2} | 5 | 70
Venting Valves and Edwards pump on gage
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | TCG | 1 | TCG:1 | XF:16IDC-VA{ES-TCG:1} | RoughingPump |  |  | 71 spare
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | BV | 1 | BV:1 | XF:16IDC-VA{ES-BV:1} | V1
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | EV | 1 | EV:1 | XF:16IDC-VA{ES-EV:1} | V2
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | EV | 2 | EV:2 | XF:16IDC-VA{ES-EV:2} | V3
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | EV | 3 | EV:3 | XF:16IDC-VA{ES-EV:3} | V4
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | EV | 4 | EV:4 | XF:16IDC-VA{ES-EV:4} | V5
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | VV | 1 | VV:1 | XF:16IDC-VA{ES-VV:1} | V1
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | VV | 2 | VV:2 | XF:16IDC-VA{ES-VV:2} | V2
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | VV | 3 | VV:3 | XF:16IDC-VA{ES-VV:3} | V3
| C |  | XF | 16IDC | XF:16IDC | VA |  | VA | ES | VV | 4 | VV:4 | XF:16IDC-VA{ES-VV:4} | V4
| B | RG:D2 | XF | 16IDB | XF:16IDB | CT |  | CT |  | VGC | 1 | VGC:1 | XF:16IDB-CT{VGC:1}
| B | RG:D2 | XF | 16IDB | XF:16IDB | CT |  | CT |  | VGC | 2 | VGC:2 | XF:16IDB-CT{VGC:2}

## PLC EPS Logic (Vac)
| Input | Signal type | Comparison | Setting | Set Point Value, mbar | AND input | AND comparison | AND value | Output | Action | Output | Action
Vacuum
Vac. Sect. | LiX ID16
A Hutch
FE | XF:16IDA-VA{DP:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | BL Vacuum Good | Reset
| XF:16IDA-VA{DP:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | BL Vacuum Good | Reset
| XF:16IDA-VA{DP:1-TCG:1} | [7 empty] | BL Vacuum Good | Reset
| XF:16IDA-VA{DP:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | BL Vacuum Good | Reset
| XF:16IDA-VA{DP:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | BL Vacuum Good | Reset
| XF:16IDA-VA{DP:1-TCG:2} | [7 empty] | BL Vacuum Good | Reset
1 | XF:16IDA-VA{DP:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{DP:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FE-GV:1} | Close
| XF:16IDA-VA{DP:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{DP:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{DP:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FE-GV:1} | Close
| XF:16IDA-VA{DP:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{DP:1-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{FE-GV:1} | Inhibit opening
| XF:16IDA-VA{DP:1-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{DP:1-GV:1} | Inhibit opening
| XF:16IDA-VA{DP:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{DP:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FE-GV:1} | Close
| XF:16IDA-VA{DP:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{DP:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{DP:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FE-GV:1} | Close
| XF:16IDA-VA{DP:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{DP:1-TCG:2} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{FE-GV:1} | Inhibit opening
| XF:16IDA-VA{DP:1-TCG:2} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{DP:1-GV:1} | Inhibit opening
2 | XF:16IDA-VA{Mir:WBM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mir:WBM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{Mir:WBM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:WBM-GV:1} | Close
| XF:16IDA-VA{Mir:WBM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mir:WBM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{DP:1-GV:1} | Close
| XF:16IDA-VA{Mir:WBM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:WBM-GV:1} | Close
| XF:16IDA-VA{Mir:WBM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{DP:1-GV:1} | Inhibit opening
| XF:16IDA-VA{Mir:WBM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mir:WBM-GV:1} | Inhibit opening
3 | XF:16IDA-VA{Mono:DMM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mono:DMM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:WBM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DMM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-IP:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mono:DMM-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:WBM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DMM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mono:DMM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:WBM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DMM-GV:1} | Close
| XF:16IDA-VA{Mono:DMM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mir:WBM-GV:1} | Inhibit opening
| XF:16IDA-VA{Mono:DMM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mono:DMM-GV:1} | Inhibit opening
4 | XF:16IDA-VA{Mono:DCM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mono:DCM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DMM-GV:1} | Close
| XF:16IDA-VA{Mono:DCM-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DCM-GV:1} | Close
| XF:16IDA-VA{Mono:DCM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{Mono:DCM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DMM-GV:1} | Close
| XF:16IDA-VA{Mono:DCM-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mono:DCM-GV:1} | Close
| XF:16IDA-VA{Mono:DCM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mono:DMM-GV:1} | Inhibit opening
| XF:16IDA-VA{Mono:DCM-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mono:DCM-GV:1} | Inhibit opening
| XF:16IDA-VA{Mono:DCM-IP:1} | Digital | = | Above setpoint | 10-5 mbar |  |  |  | XF:16IDA-UT{Cryo:1}Cmd:Run-Cmd | Stop
| XF:16IDA-VA{Cryo:1-TCG:1} | Digital | = | Above setpoint | 1 |  |  |  | XF:16IDA-UT{Cryo:1}Cmd:Run-Cmd | Stop
5 | XF:16IDA-VA{FS:3-IP:1} | Digital | = | Above setpoint |  |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDA-VA{FS:3-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | XF:16IDA-VA{Mono:DCM-GV:1} | Close
| XF:16IDA-VA{FS:3-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | XF:16IDA-VA{FS:3-GV:1} | Close
| XF:16IDA-VA{FS:3-CCG:1} | Digital | = | Above setpoint |  |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDA-VA{FS:3-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | XF:16IDA-VA{Mono:DCM-GV:1} | Close
| XF:16IDA-VA{FS:3-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | XF:16IDA-VA{FS:3-GV:1} | Close
| XF:16IDA-VA{FS:3-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mono:DCM-GV:1} | Inhibit opening
| XF:16IDA-VA{FS:3-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{FS:3-GV:1} | Inhibit opening
6 | XF:16IDA-VA{Mir:KB-IP:1} | Digital | = | Above setpoint |  |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDA-VA{Mir:KB-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FS:3-GV:1} | Close
| XF:16IDA-VA{Mir:KB-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDA-VA{Mir:KB-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDA-VA{Mir:KB-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{FS:3-GV:1} | Close
| XF:16IDA-VA{Mir:KB-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDA-VA{Mir:KB-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{FS:3-GV:1} | Inhibit opening
| XF:16IDA-VA{Mir:KB-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mir:KB-GV:1} | Inhibit opening
KB Mirror power supplys must be shut of if preassure is above 10 -5 mbar
| XF:16IDA-VA{Mir:KB-CCG:1} | Digital | = | Above setpoint | 10-7 mbar |  |  |  | XF:16IDA-VA{Mir:KB-Trip:HV} | Kill KB power supply
| XF:16IDA-VA{Mir:KB-IP:1} | Digital | = | Above setpoint | 10-7 mbar |  |  |  | XF:16IDA-VA{Mir:KB-Trip:HV} | Kill KB power supply | monochromatic beam
7 | XF:16IDA-VA{BT:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close
| XF:16IDA-VA{BT:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDA-VA{BT:1-IP:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDB-VA{BT:1-GV:1} | Close
| XF:16IDA-VA{BT:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDA-VA{BT:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDA-VA{BT:1-CCG:1} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDB-VA{BT:1-GV:1} | Close
| XF:16IDA-VA{Mir:KB-TCG:1} | Digital | = | Above setpoint |  |  |  |  | XF:16IDA-VA{FS:3-GV:1} | Inhibit opening
| XF:16IDA-VA{Mir:KB-TCG:1} | Digital | = | Above setpoint |  |  |  |  | XF:16IDA-VA{Mir:KB-GV:1} | Inhibit opening
PIPE
B Hutch
| XF:16IDB-VA{BT:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDB-VA{BT:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDB-VA{BT:1-IP:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDB-VA{BT:1-GV:1} | Close
| XF:16IDB-VA{BT:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar |  |  |  | FE Shutter | Close | monochromatic beam
| XF:16IDB-VA{BT:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDA-VA{Mir:KB-GV:1} | Close
| XF:16IDB-VA{BT:1-CCG:2} | Digital | = | Above setpoint | 10 -7 mbar | Beam in Front End | = | FALSE | XF:16IDB-VA{BT:1-GV:1} | Close
| XF:16IDB-VA{BT:1-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDA-VA{Mir:KB-GV:1} | Inhibit opening
| XF:16IDB-VA{BT:1-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar |  |  |  | XF:16IDB-VA{BT:1-GV:1} | Inhibit opening
8
Auto Evacuation and Venting
Init; New EM is installed;
PUMP IS ON , if CCG shows good(10-3) vacuum
| XF:16IDC-VA{ES-TCG:1} | Digital | = | Above setpoint | 10 -3 mbar
[9 empty] | XF:16IDB-VA{BT:1-GV:1} | Close | ES_GV0
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Close | ES_GV1
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Close | ES_GV2
[9 empty] | XF:16IDC-VA{EM-GV:1} | Close | ES_GV3
[10 empty] | Open | V1
[10 empty] | Open | V2
[10 empty] | Open | V3
[10 empty] | Open | V4
[10 empty] | Open | V5
| No EM signal; Maintanace; Chamber(s) venting | [8 empty] | `
[9 empty] | XF:16IDB-VA{BT:1-GV:1} | Close | ES_GV0
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Close | ES_GV1
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Close | ES_GV2
[9 empty] | XF:16IDC-VA{EM-GV:1} | Close | ES_GV3
[10 empty] | Open | V1
[10 empty] | Close | V2
[10 empty] | Close | V3
[10 empty] | Close | V4
[10 empty] | Close | V5
| Experiment | [7 empty] | XF:16IDB-VA{BT:1-GV:1} | Open | ES_GV0
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Open | ES_GV1
[9 empty] | XF:16IDC-VA{Chm:SF-GV:1} | Open | ES_GV2
[9 empty] | XF:16IDC-VA{EM-GV:1} | Open | ES_GV3
[10 empty] | Close | V1
[10 empty] | Open | V2
[10 empty] | Open | V3
[10 empty] | Open | V4
[10 empty] | Open | V5
Condition1
| XF:16IDB-VA{BT:1-IP:2} | Digital | = | Below setpoint 10-8 |  | Vacuum in EM section | = | Below setpoint 10-2 | and -sel (user selection)
EM type A
EM type B
EM type C
INIT
EXPERIMENT
MAINTANACE
Condition 1

## Auto Evacuation Venting
GV0
GV1
GV2
GV3
GV4
State Name | Input | Input, short name | Signal type | Comparison | Setting | Set Point Value | AND input | AND comparison | AND value | Output | Action | Output | Action
Init0 |  | GV0 |  | = | closed
|  | GV1 |  | = | closed
|  | GV2 |  | = | closed
|  | GV3 |  | = | closed
|  | V1 |  | = | open
|  | V2 |  | = | closed
|  | V3 |  | = | closed
|  | V4 |  | = | ?????
|  | V5 |  | = | closed
Init1
| Input | Signal type |  | Setting | AND input |  | AND comparison | AND value | Output | Action | Output | Action
| vacuum reading on the beamline | Digital |  | Above setpoint (10-8 mbar) | GV0 |  | = | closed | INIT

## Vacuum Controllers A
Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12 | Column13 | Column14 | Column15 | Column16
16 ID-A
| VAC SEC 1 |  | VAC SEC 2 |  | VAC SEC 3 |  |  |  | VAC SEC 4 |  | VAC SEC 5 |  | VAC SEC 6
| 27.4 |  | 29.1 |  | 30.7 |  |  |  | 32.7 |  | 36.3 |  | 37.7
| VGC-1A |  | VGC-2A |  | VGC-2B |  |  |  | VGC-3B |  | VGC-4A |  | *FUTURE
| 38-CCG HV |  | 41-CCG HV |  | 36-CCG HV |  |  |  | 8-CCG HV |  | 39-CCG HV |  | *37-CCG HV
| 47-CCG SIG |  | 50-CCG SIG |  | 45-CCG SIG |  |  |  | 43-CCG SIG |  | 48-CCG SIG |  | *46-CCG SIG
| 99-PIRANI |  | 107-PIRANI |  | 103-PIRANI |  |  |  | 101-PIRANI |  | 105-PIRANI |  | *106-PIRANI
| 28.2 |  |  |  | 31.5 | [7 empty] | 39
| VGC-1B |  |  |  | VGC-3A | [7 empty] | VGC-4B
| 7-CCG HV |  |  |  | 109-CCG HV |  |  |  | <-KB mir Vac sec 6 |  |  |  | 40-CCG HV
| 42-CCG SIG |  |  |  | 44-CCG SIG |  |  |  | <-KB mir Vac sec 6 |  |  |  | 49-CCG SIG
| 100-PIRANI |  |  |  | 104-PIRANI |  |  |  | <-DCM CRIO |  | KB mir Vac sec 6-> |  | 102-PIRANI
| IPC-1 |  | IPC-2 |  | IPC-2 |  |  |  | IPC-3 |  | IPC-4 |  | IPC-4
| 74-IP |  | 77-IP |  | 72-IP |  |  |  | 70-IP |  | 75-IP |  | 71-IP
69-IP
[5 empty] | IPC-3 | [7 empty] | IPC-5
[5 empty] | IP-73 | [7 empty] | 76-IP
*FUTURE CABLES / NOT CONNECTED

## Motion axis summary
Index | Component | Component ID | Full component name | Motion axis | Axis ID | Axis name | Full axis name | Motor | [9 empty] | Encoder |  |  |  |  | Switches |  |  | Control
[8 empty] | Type | Manufacturer | Model | Resolution | Current | Phase resistance (ohms) | Phase inductance (mH) | Gearbox | Microstep ratio | Overall resolution (EGU/full step) | Type | Manufacturer | Readhead Model | Interpolator Model | Resolution (micron) | Limit | Reference | Overtravel | Rack | Controller type | Controller | Axis | Motor cable number | Encoder cable number | IOC Name | EGU | PREC | MRES | ALIAS | Homing PLC
1 | LiX FOE Optics Hutch A
2 | Beam-viewing screen #1
3 | Fixed Mask
4 | White Beam Mirror | Mir:WBM | XF:16IDA-OP{Mir:WBM} | X | X | Ax:X | XF:16IDA-OP{Mir:WBM-Ax:X}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 1 | 96 | 66 | mc1 | mm | 4 | 0.0001 |  | 8
5 | White Beam Mirror | Mir:WBM | XF:16IDA-OP{Mir:WBM} | Y | Y | Ax:Y | XF:16IDA-OP{Mir:WBM-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 2 | 97 | 67 | mc1 | mm | 4 | 0.0001 |  | 9
6 | White Beam Mirror | Mir:WBM | XF:16IDA-OP{Mir:WBM} | Pitch | P | Ax:P | XF:16IDA-OP{Mir:WBM-Ax:P}Mtr | [18 empty] | RG:A1 | GBLV | 1 | 3 | 98 | 68 | mc1 | mrad | 6 | 1.1e-05 |  | 10
7 | White Beam Mirror | Mir:WBM | XF:16IDA-OP{Mir:WBM} | Fine Pitch; Out of Scope | PF | Ax:PF | XF:16IDA-OP{Mir:WBM-Ax:PF}Mtr | [18 empty] | RG:A1 | piezo TBD |  |  | NA
8 | Beam-viewing screen #2 | FS:2 | XF:16IDA-BI{FS:2} |  | Y | Ax:Y | XF:16IDA-BI{FS:2-Ax:Y}Mtr | [18 empty] | RG:A1 |  |  |  | NA
9 | Dual Multilayer Monochromator |  |  | theta1; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 4 | 114 | 118 |  | tbd
10 | Dual Multilayer Monochromator |  |  | X1; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 5 | 115 | 119 |  | tbd
11 | Dual Multilayer Monochromator |  |  | theta2; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 6 | 116 | 120 |  | tbd
12 | Dual Multilayer Monochromator |  |  | X2; Out of Scope | [21 empty] | RG:A1 | GBLV | 1 | 7 | 117 | 121 |  | tbd
13 | Dual Multilayer Monochromator |  |  | roll 2; Out of Scope | [21 empty] | RG:A1 | PicoMotor MCS |  |  | NA
14 | Dual Multilayer Monochromator |  |  | fine pitch 2; Out of Scope | [21 empty] | RG:A1 | piezoTBD |  |  | NA
15 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | theta, Bragg | Bragg | Ax:Bragg | XF:16IDA-OP{Mono:DCM-Ax:Bragg}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 1 | 78 | 51 | mc2 | deg | 8 | 6.26e-06 |  | 8
16 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | X | X | Ax:X | XF:16IDA-OP{Mono:DCM-Ax:X}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 2 | 79 | 52 | mc2 | mm | 3 | 0.001 |  | 9
17 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | DCM.Offset2 (FE.Y) | Of2 | Ax:Of2 | XF:16IDA-OP{Mono:DCM-Ax:Of2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 3 | 80 | 53 | mc2 | mm | 4 | 0.0002 |  | 10
18 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | DCM.Pitch2 | P2 | Ax:P2 | XF:16IDA-OP{Mono:DCM-Ax:P2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 4 | 81 | 54 | mc2 | mm | 6 | 1e-06 |  | 11 | no enc
19 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | DCM.Roll2 | R2 | Ax:R2 | XF:16IDA-OP{Mono:DCM-Ax:R2}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 5 | 82 | 55 | mc2 | mm | 6 | 2e-06 |  | 12
20 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | DCM.finePitch piezo | PF2 | Ax:PF2 | XF:16IDA-OP{Mono:DCM-Ax:PF2}Mtr | [18 empty] | RG:A1 | PI? |  |  | NA
21 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | CCM.finePitch stepper | CCM_PF | Ax:CCM_PF | XF:16IDA-OP{Mono:DCM-Ax:CCM_PF}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 6 | 83 | NA | mc2 | mm | 4 | 0.0001 |  | 13
22 | Si(111) Monochromator DCM | Mono:DCM | XF:16IDA-OP{Mono:DCM} | DCM.finePitch readback only | PF_RDBK | Ax:PF_RDBK | XF:16IDA-OP{Mono:DCM-Ax:PF_RDBK}Mtr | [18 empty] | RG:A1 | GBLV | 2 | 7 | NA | 56 | mc2 | mm | 6 | 1e-06 |  | 14 | encoder only; this encoder is on 4
23 | Beam-viewing screen #3 | FS:3 | XF:16IDA-BI{FS:3} | Y | Y | Ax:Y | XF:16IDA-BI{FS:3-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 3 | 95 | NA | mc4 | mm | 6 | 0.05 |  | 11
24 | Slits | Slt:1 | XF:16IDA-OP{Slt:1} | X-inboard | I | Ax:I | XF:16IDA-OP{Slt:1-Ax:I}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 5 | 91 | NA | mc4 | mm | 6 | 0.05 |  | 12
25 | Slits | Slt:1 | XF:16IDA-OP{Slt:1} | X-outboard | O | Ax:O | XF:16IDA-OP{Slt:1-Ax:O}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 6 | 92 | NA | mc4 | mm | 6 | 0.05 |  | 13
26 | Slits | Slt:1 | XF:16IDA-OP{Slt:1} | Y-top | T | Ax:T | XF:16IDA-OP{Slt:1-Ax:T}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 7 | 93 | NA | mc4 | mm | 6 | 0.05 |  | 14
27 | Slits | Slt:1 | XF:16IDA-OP{Slt:1} | Y-bottom | B | Ax:B | XF:16IDA-OP{Slt:1-Ax:B}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 8 | 94 | NA | mc4 | mm | 6 | 0.05 |  | 15
28 | KB mirror system Horizontal Focusing Mirror | Mir:KBH | XF:16IDA-OP{Mir:KBH} | HFMX-upstream | XU | Ax:XU | XF:16IDA-OP{Mir:KBH-Ax:XU}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 1 | 84 | 57 | mc3 | mm | 4 | 0.00024582 |  | 8
29 | KB mirror system Horizontal Focusing Mirror | Mir:KBH | XF:16IDA-OP{Mir:KBH} | HFMX-downstream | XD | Ax:XD | XF:16IDA-OP{Mir:KBH-Ax:XD}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 2 | 85 | 58 | mc3 | mm | 4 | 0.00024582 |  | 9
30 | KB mirror system Horizontal Focusing Mirror | Mir:KBH | XF:16IDA-OP{Mir:KBH} | HFMY-upstream | YU | Ax:YU | XF:16IDA-OP{Mir:KBH-Ax:YU}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 3 | 86 | 59 | mc3 | mm | 4 | 0.00049164 |  | 10
31 | KB mirror system Horizontal Focusing Mirror | Mir:KBH | XF:16IDA-OP{Mir:KBH} | HFMY-downstream | YD | Ax:YD | XF:16IDA-OP{Mir:KBH-Ax:YD}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 4 | 87 | 60 | mc3 | mm | 4 | 0.00049164 |  | 11
32 | KB mirror system Vertical Focusing Mirror | Mir:KBV | XF:16IDA-OP{Mir:KBV} | VFMX one horizontal translation | X | Ax:X | XF:16IDA-OP{Mir:KBV-Ax:X}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 5 | 88 | 61 | mc3 | mm | 4 | 0.00024582 |  | 12
33 | KB mirror system Vertical Focusing Mirror | Mir:KBV | XF:16IDA-OP{Mir:KBV} | VFMY-upstream | YU | Ax:YU | XF:16IDA-OP{Mir:KBV-Ax:YU}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 6 | 89 | 62 | mc3 | mm | 4 | 0.00049164 |  | 13
34 | KB mirror system Vertical Focusing Mirror | Mir:KBV | XF:16IDA-OP{Mir:KBV} | VFMY-downstream | YD | Ax:YD | XF:16IDA-OP{Mir:KBV-Ax:YD}Mtr | [18 empty] | RG:A1 | GBLV | 3 | 7 | 90 | 63 | mc3 | mm | 4 | 0.00049164 |  | 14
35 | KB mirror system Horizontal Focusing Mirror | Mir:KBH | XF:16IDA-OP{Mir:KBH} | Fine Pitch Adjustment HFM RDBK | PF_RDBK | Ax:PF_RDBK | XF:16IDA-OP{Mir:KBH-Ax:PF_RDBK}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 1 | NA | 64 | mc4 | mm | 6 | 0.0005 |  | 8
37 | KB mirror system Vertical Focusing Mirror | Mir:KBV | XF:16IDA-OP{Mir:KBV} | Fine Pitch Adjustment VFM RDBK | PF_RDBK | Ax:PF_RDBK | XF:16IDA-OP{Mir:KBV-Ax:PF_RDBK}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 2 | NA | 65 | mc4 | mm | 6 | 0.0005 |  | 9
39 | Beam-viewing screen #4 | FS:4 | XF:16IDA-BI{FS:4} | Y | Y | Ax:Y | XF:16IDA-BI{FS:4-Ax:Y}Mtr | [18 empty] | RG:A1 | GBLV | 4 | 4 | 108 | NA | mc4 | mm | 6 | 0.05 |  | 10
40
41
42
43 | Secondry Source SOE Hutch B
44 | Beam Position Monitor (BPM) | BPM:1 | XF:16IDB-BI{BPM:1} | X | X | Ax:X | XF:16IDB-BI{BPM:1-Ax:X}Mtr | [18 empty] | RG:B1 | SmartAct MCS |  |  |  |  | mc | um | 3 | 1
44 | Beam Position Monitor (BPM) | BPM:1 | XF:16IDB-BI{BPM:1} | Y | Y | Ax:Y | XF:16IDB-BI{BPM:1-Ax:Y}Mtr | [18 empty] | RG:B1 | SmartAct MCS |  |  |  |  | mc | um | 3 | 1
45 | Secondary Source Aperture(SSA) | Slt:SSA | XF:16IDB-OP{Slt:SSA} | X | X | Ax:X | XF:16IDB-OP{Slt:SSA-Ax:X}Mtr | [18 empty] | RG:B1 | PiezoSysJena |  |  |  | 0 | mc | um | 3 | 1
46 | Secondary Source Aperture(SSA) | Slt:SSA | XF:16IDB-OP{Slt:SSA} | Y | Y | Ax:Y | XF:16IDB-OP{Slt:SSA-Ax:Y}Mtr | [18 empty] | RG:B1 | PiezoSysJena |  |  |  | 0 | mc | um | 3 | 1
47 | Attenuator | Fltr:Atn | XF:16IDB-OP{Fltr:Atn} | absorber set #1 | X1 | Ax:X1 | XF:16IDB-OP{Fltr:Atn-Ax:X1}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 5 | B-5 | B-6 | mc5 | um | 3 | 1 |  | 12
48 | Attenuator | Fltr:Atn | XF:16IDB-OP{Fltr:Atn} | absorber set #2 | X2 | Ax:X2 | XF:16IDB-OP{Fltr:Atn-Ax:X2}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 6 | B-9 | B-7 | mc5 | um | 3 | 1 |  | 13
49 | Attenuator | Fltr:Atn | XF:16IDB-OP{Fltr:Atn} | absorber set #3 | X3 | Ax:X3 | XF:16IDB-OP{Fltr:Atn-Ax:X3}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 7 | B-10 | B-8 | mc5 | um | 3 | 1 |  | 14
50 | Visual Beam Monitor (VBM) | FS:VBM | XF:16IDB-BI{FS:VBM} | focus | F | Ax:F | XF:16IDB-BI{FS:VBM-Ax:F}Mtr | [18 empty] | RG:B1 | PicoMotor MCS |  |  |  |  | mc | um | 3 | 1
51 | Visual Beam Monitor (VBM) | FS:VBM | XF:16IDB-BI{FS:VBM} | zoom | Zm | Ax:Zm | XF:16IDB-BI{FS:VBM-Ax:Zm}Mtr | [18 empty] | RG:B1 | PicoMotor MCS |  |  |  |  | mc | um | 3 | 1
52 | Shutter | Sh | XF:16IDB-OP{Sh} | open | X | Ax:X | XF:16IDB-OP{Sh-Ax:X}Mtr | [18 empty] | RG:B1 | PLC |  |  |  |  | mc | um | 3 | 1
53 | Alternative SSA | Slt:aSSA | XF:16IDB-OP{Slt:aSSA} | X | X | Ax:X | XF:16IDB-OP{Slt:aSSA-Ax:X}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 1 | B-11 |  | mc5 | um | 3 | 1.2393e-05 |  | 8
54 | Alternative SSA | Slt:aSSA | XF:16IDB-OP{Slt:aSSA} | dX | dX | Ax:dX | XF:16IDB-OP{Slt:aSSA-Ax:dX}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 2 | B-12 |  | mc5 | um | 3 | 3.1758e-05 |  | 9
55 | Alternative SSA | Slt:aSSA | XF:16IDB-OP{Slt:aSSA} | Y | Y | Ax:Y | XF:16IDB-OP{Slt:aSSA-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 3 | B-22 |  | mc5 | um | 3 | 1.2393e-05 |  | 10
56 | Alternative SSA | Slt:aSSA | XF:16IDB-OP{Slt:aSSA} | dY | dY | Ax:dY | XF:16IDB-OP{Slt:aSSA-Ax:dY}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 4 | B-23 |  | mc5 | um | 3 | 3.1758e-05 |  | 11
57
58
59 | Experimental End Station Enclosure EESE Hutch C
60 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | Z | Z | Ax:Z | XF:16IDC-OP{CRL-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 1 | C-7 |  | mc9 | um | 3 | 1
61 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | Z1 | Z1 | Ax:Z1 | XF:16IDC-OP{CRL-Ax:Z1}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 2 | C-8 |  | mc9 | um | 3 | 1
62 | Harmonic Regection Mirror HRM1 | Mir:HRM1 | XF:16IDC-OP{Mir:HRM1} | Y | Y | Ax:Y | XF:16IDC-OP{Mir:HRM1-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 2 | C-20 |  | mc6 | um | 3 | 1
63 | Harmonic Regection Mirror HRM1 | Mir:HRM1 | XF:16IDC-OP{Mir:HRM1} | theta | Th | Ax:Th | XF:16IDC-OP{Mir:HRM1-Ax:Th}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 1 | C-19 |  | mc6 | um | 3 | 1
64 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | X Upstream | UX | Ax:UX | XF:16IDC-OP{CRL-Ax:UX}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 5 | C-11 |  | mc6 | um | 3 | 1
65 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | Y Upstream | UY | Ax:UY | XF:16IDC-OP{CRL-Ax:UY}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 6 | C-12 |  | mc6 | um | 3 | 1
66 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | X Downstream | DX | Ax:DX | XF:16IDC-OP{CRL-Ax:DX}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 7 | C-13 |  | mc6 | um | 3 | 1
67 | Transfocator CRLs | CRL | XF:16IDC-OP{CRL} | Y Downstream | DY | Ax:DY | XF:16IDC-OP{CRL-Ax:DY}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 8 | C-14 |  | mc6 | um | 3 | 1
68 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-OP{Mir:HRM2} | Y | Y | Ax:Y | XF:16IDC-OP{Mir:HRM2-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 6 | 3 | C-21 |  | mc6 | um | 3 | 1
69 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-OP{Mir:HRM2} | theta, tilt | Th | Ax:Th | XF:16IDC-OP{Mir:HRM2-Ax:Th}Mtr | [18 empty] | RG:B1 | GBLV | 5 | 8 | C-22 | C-23 | mc5 | rad | 3 | 1 |  | 15
70 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-OP{Mir:HRM2} | Upstream Bend; with LVDT | BU | Ax:BU | XF:16IDC-OP{Mir:HRM2-Ax:BU}Mtr | [18 empty] | RG:C2 | Picomotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
71 | Harmonic Regection Mirror HRM2 | Mir:HRM2 | XF:16IDC-OP{Mir:HRM2} | Downstream Bend; With LVDT | BD | Ax:BD | XF:16IDC-OP{Mir:HRM2-Ax:BD}Mtr | [18 empty] | RG:C2 | Picomotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
72 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-OP{Slt:DDA} | X | X | Ax:X | XF:16IDC-OP{Slt:DDA-Ax:X}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 1 | C-15 |  | mc7 | um | 3 | 1
73 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-OP{Slt:DDA} | Y | Y | Ax:Y | XF:16IDC-OP{Slt:DDA-Ax:Y}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 2 | C-16 |  | mc7 | um | 3 | 1
74 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-OP{Slt:DDA} | dX | dX | Ax:dX | XF:16IDC-OP{Slt:DDA-Ax:dX}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 3 | C-17 |  | mc7 | um | 3 | 1
75 | Divergence Defining Aperture (DDA) | Slt:DDA | XF:16IDC-OP{Slt:DDA} | dY | dY | Ax:dY | XF:16IDC-OP{Slt:DDA-Ax:dY}Mtr | [18 empty] | RG:B1 | GBLV | 7 | 4 | C-18 |  | mc7 | um | 3 | 1
76 | Beam Position Monitor (BPM) | BPM:2 | XF:16IDC-BI{BPM:2} | Y | Y | Ax:Y | XF:16IDC-BI{BPM:2-Ax:Y}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
77 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-OP{Slt:G1} | X | X | Ax:X | XF:16IDC-OP{Slt:G1-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 1 | C-9 |  | mc10 | um | 3 | 1
78 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-OP{Slt:G1} | dX | dX | Ax:dX | XF:16IDC-OP{Slt:G1-Ax:dX}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 2 | C-10 |  | mc10 | um | 3 | 1
79 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-OP{Slt:G1} | Y | Y | Ax:Y | XF:16IDC-OP{Slt:G1-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 3 | C-25 |  | mc10 | um | 3 | 1
80 | Guard Slits 1 in Air | Slt:G1 | XF:16IDC-OP{Slt:G1} | dY | dY | Ax:dY | XF:16IDC-OP{Slt:G1-Ax:dY}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 4 | C-26 |  | mc10 | um | 3 | 1
81 | In-air transmission measurement module
82 | Coarse-resolution scanning | Stg:ScanC | XF:16IDC-ES:InAir{Stg:ScanC} | scan X | X | Ax:X | XF:16IDC-ES:InAir{Stg:ScanC-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 1 | C-34 |  | mc8 | um | 3 | 1
83 | Coarse-resolution scanning | Stg:ScanC | XF:16IDC-ES:InAir{Stg:ScanC} | scan Y | Y | Ax:Y | XF:16IDC-ES:InAir{Stg:ScanC-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 2 | C-35 |  | mc8 | um | 3 | 1
84 | Fine-resolution scanning | Stg:ScanF | XF:16IDC-ES:InAir{Stg:ScanF} | scan X | X | Ax:X | XF:16IDC-ES:InAir{Stg:ScanF-Ax:X}Mtr | [18 empty] | RG:C2 | PI MCS | PI |  |  |  | mcPI | um | 3 | 1
85 | Fine-resolution scanning | Stg:ScanF | XF:16IDC-ES:InAir{Stg:ScanF} | scan Y | Y | Ax:Y | XF:16IDC-ES:InAir{Stg:ScanF-Ax:Y}Mtr | [18 empty] | RG:C2 | PI MCS | PI |  |  |  | mcPI | um | 3 | 1
86 | Fine-resolution scanning | Stg:ScanF | XF:16IDC-ES:InAir{Stg:ScanF} | azimuthal | Rot | Ax:Rot | XF:16IDC-ES:InAir{Stg:ScanF-Ax:Rot}Mtr | [18 empty] | RG:C2 | GBLV | 6 | 4 | C-24 |  | mc6 | um | 3 | 1
87 | Microscope | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1} | X | X | Ax:X | XF:16IDC-ES:InAir{Mscp:1-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 4 | C-31 |  | mc8 | um | 3 | 1
88 | Microscope | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1} | Y | Y | Ax:Y | XF:16IDC-ES:InAir{Mscp:1-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 5 | C-32 |  | mc8 | um | 3 | 1
89 | Microscope | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1} | focus | F | Ax:F | XF:16IDC-ES:InAir{Mscp:1-Ax:F}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 8 | C-33 |  | mc10 | um | 3 | 1
90 | Microscope | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1} | polarizer/analyze | Pol | Ax:Pol | XF:16IDC-ES:InAir{Mscp:1-Ax:Pol }Mtr | [18 empty] | RG:B1 | GBLV | 7 | 7 | C-106 |  | mc7 | um | 3 | 1
91 | Microscope | Mscp:1 | XF:16IDC-ES:InAir{Mscp:1} | zoom | Zm | Ax:Zm | XF:16IDC-ES:InAir{Mscp:1-Ax:Zm}Mtr | [19 empty] | Fetura RS-232
92 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:InAir{Slt:G2} | X | X | Ax:X | XF:16IDC-ES:InAir{Slt:G2-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 1 | C-27 |  | mc11 | um | 3 | 1
93 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:InAir{Slt:G2} | dX | dX | Ax:dX | XF:16IDC-ES:InAir{Slt:G2-Ax:dX}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 2 | C-28 |  | mc11 | um | 3 | 1
94 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:InAir{Slt:G2} | Y | Y | Ax:Y | XF:16IDC-ES:InAir{Slt:G2-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 3 | C-29 |  | mc11 | um | 3 | 1
95 | Guard Slits 2 in Air | Slt:G2 | XF:16IDC-ES:InAir{Slt:G2} | dY | dY | Ax:dY | XF:16IDC-ES:InAir{Slt:G2-Ax:dY}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 4 | C-30 |  | mc11 | um | 3 | 1
96 | In-vaccum GISAXS/GID module
97 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | X | X | Ax:X | XF:16IDC-ES:GI{Stg:Hex-Ax:X}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
98 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | Y | Y | Ax:Y | XF:16IDC-ES:GI{Stg:Hex-Ax:Y}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
99 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | Z | Z | Ax:Z | XF:16IDC-ES:GI{Stg:Hex-Ax:Z}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
100 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | U | U | Ax:U | XF:16IDC-ES:GI{Stg:Hex-Ax:U}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
101 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | V | V | Ax:V | XF:16IDC-ES:GI{Stg:Hex-Ax:V}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
102 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | W | W | Ax:W | XF:16IDC-ES:GI{Stg:Hex-Ax:W}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
103 | Hexapod | Stg:Hex | XF:16IDC-ES:GI{Stg:Hex} | phi | phi | Ax:phi | XF:16IDC-ES:GI{Stg:Hex-Ax:phi}Mtr | [18 empty] | RG:C2 | micos MCS |  |  |  |  | mc | um | 3 | 1
104 | Guard Slits 2 in VAC | Slt:G2 | XF:16IDC-ES:GI{Slt:G2} | X | X | Ax:X | XF:16IDC-ES:GI{Slt:G2-Ax:X}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
105 | Guard Slits 2 in VAC | Slt:G2 | XF:16IDC-ES:GI{Slt:G2} | dX | dX | Ax:dX | XF:16IDC-ES:GI{Slt:G2-Ax:dX}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
106 | Sample handling
107 | Solution sample handler | SM:Sol | XF:16IDC-ES:Soln{SM:Sol} | X | X | Ax:X | XF:16IDC-ES:Soln{SM:Sol-Ax:X}Mtr | [18 empty] | RG:C2 | TBD MCS/GBLV | TBD |  |  |  | mcTBD | um | 3 | 1
108 | Solution sample handler | SM:Sol | XF:16IDC-ES:Soln{SM:Sol} | Y | Y | Ax:Y | XF:16IDC-ES:Soln{SM:Sol-Ax:Y}Mtr | [18 empty] | RG:C2 | TBD MCS/GBLV | TBD |  |  |  | mcTBD | um | 3 | 1
109 | Staubli TX40 robot | SM:Staubli | XF:16IDC-ES{SM:Staubli} | routins will be called via PVs | TBD | Ax:TBD | XF:16IDC-ES{SM:Staubli-Ax:TBD }Mtr | [18 empty] | RG:C2 | Staubli MCS | Staubli |  |  |  | mcStaubli | um | 3 | 1
110 | Detector system
111 | WAXS Detector Positioning Stage 1 | Stg:WAXS1 | XF:16IDC-ES{Stg:WAXS1} | X | X | Ax:X | XF:16IDC-ES{Stg:WAXS1-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 3 | C-36 |  | mc9 | um | 3 | 1
112 | WAXS Detector Positioning Stage 1 | Stg:WAXS1 | XF:16IDC-ES{Stg:WAXS1} | Y | Y | Ax:Y | XF:16IDC-ES{Stg:WAXS1-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 4 | C-37 |  | mc9 | um | 3 | 1
113 | WAXS Detector Positioning Stage 1 | Stg:WAXS1 | XF:16IDC-ES{Stg:WAXS1} | Z | Z | Ax:Z | XF:16IDC-ES{Stg:WAXS1-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 5 | C-38 |  | mc9 | um | 3 | 1
114 | WAXS Detector Positioning Stage 2 | Stg:WAXS2 | XF:16IDC-ES{Stg:WAXS2} | X | X | Ax:X | XF:16IDC-ES{Stg:WAXS2-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 6 | C-39 |  | mc9 | um | 3 | 1
115 | WAXS Detector Positioning Stage 2 | Stg:WAXS2 | XF:16IDC-ES{Stg:WAXS2} | Y | Y | Ax:Y | XF:16IDC-ES{Stg:WAXS2-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 7 | C-40 |  | mc9 | um | 3 | 1
116 | WAXS Detector Positioning Stage 2 | Stg:WAXS2 | XF:16IDC-ES{Stg:WAXS2} | Z | Z | Ax:Z | XF:16IDC-ES{Stg:WAXS2-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 9 | 8 | C-41 |  | mc9 | um | 3 | 1
117 | WAXS beamstop | BS:WAXS | XF:16IDC-ES{BS:WAXS} | X | X | Ax:X | XF:16IDC-ES{BS:WAXS-Ax:X}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
118 | WAXS beamstop | BS:WAXS | XF:16IDC-ES{BS:WAXS} | Y | Y | Ax:Y | XF:16IDC-ES{BS:WAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
119 | WAXS beamstop | BS:WAXS | XF:16IDC-ES{BS:WAXS} | Z | Z | Ax:Z | XF:16IDC-ES{BS:WAXS-Ax:Z}Mtr | [18 empty] | RG:C2 | SmarAct MCS | SmartAct |  |  |  | mcSmartAct | um | 3 | 1
120 | SAXS Detector Positioning Stage | Stg:SAXS | XF:16IDC-ES{Stg:SAXS} | X | X | Ax:X | XF:16IDC-ES{Stg:SAXS-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 6 | C-48 |  | mc8 | um | 3 | 1
121 | SAXS Detector Positioning Stage | Stg:SAXS | XF:16IDC-ES{Stg:SAXS} | Y | Y | Ax:Y | XF:16IDC-ES{Stg:SAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 7 | C-49 |  | mc8 | um | 3 | 1
122 | SAXS Detector Positioning Stage | Stg:SAXS | XF:16IDC-ES{Stg:SAXS} | Z | Z | Ax:Z | XF:16IDC-ES{Stg:SAXS-Ax:Z}Mtr | [18 empty] | RG:C2 | GBLV | 8 | 8 | C-50 | C-45 C-46 | mc8 | um | 3 | 1
123 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | X | X | Ax:X | XF:16IDC-ES{BS:SAXS-Ax:X}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 5 | C-51 |  | mc11 | um | 3 | 1
124 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | Y | Y | Ax:Y | XF:16IDC-ES{BS:SAXS-Ax:Y}Mtr | [18 empty] | RG:C2 | GBLV | 11 | 6 | C-52 |  | mc11 | um | 3 | 1
125 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | R1 | R1 | Ax:R1 | XF:16IDC-ES{BS:SAXS-Ax:R1}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 5 | C-65 |  | mc10 | um | 3 | 1
126 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | R2 | R2 | Ax:R2 | XF:16IDC-ES{BS:SAXS-Ax:R2}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 6 | C-66 |  | mc10 | um | 3 | 1
127 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | R3 | R3 | Ax:R3 | XF:16IDC-ES{BS:SAXS-Ax:R3}Mtr | [18 empty] | RG:C2 | GBLV | 10 | 7 | C-67 |  | mc10 | um | 3 | 1
128 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | tilt1 | T1 | Ax:T1 | XF:16IDC-ES{BS:SAXS-Ax:T1}Mtr | [18 empty] | RG:C2 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
129 | SAXS beam stop | BS:SAXS | XF:16IDC-ES{BS:SAXS} | tilt2 | T2 | Ax:T2 | XF:16IDC-ES{BS:SAXS-Ax:T2}Mtr | [18 empty] | RG:C2 | PicoMotor MCS | Picomotor |  |  |  | mcPicomotor | um | 3 | 1
Note on B97: This is part of the GISAXS/GID module

## GBLV Motor Controllers
Beamline: | LiX
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
Controler | Controller #1 | Controller #2 | Controller #3 | Controller #4 | Controller #5 | Controller #6 | Controller #7 | Controller #8 | Controller #9 | Controller #10 | Controller #11 | Controller #12 | Controller #13
Location | Rack A-1 | Rack A-1 | Rack A-1 | Rack A-1 | Rack B-1 | Rack B-1 | Rack B-1 | Rack C-2 | Rack C-2 | Rack C-2 | Rack C-2 | Spare | Spare
Model # | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000 | BHB8-C0-112-00000000 | BHB8-C0-442-00000000 | BHB8-C0-112-00000000
Channels' description | 8x5A | 8x5A | 8x5A | 8x5A | 4x1A, 4x0.25A | 8x5A | 4x1A, 4x0.25A | 8x5A | 8x5A | 4x1A, 4x0.25A | 4x1A, 4x0.25A | 8x5A | 4x1A, 4x0.25A
Channel 1 | White Beam Mirror X | Monochromator theta | KB mirror sys HFMX-upstream | KB HFM fine pitch (readback only) | Alternative SSA X (450mA) | HRM1 theta | Divergence-defining aperture X (450mA) | In-air measurement Scan X | Secondary Focusing working distance Z | Slits Guard SG1 X (450mA) | Slits Guard SG2 X (450mA)
Channel 2 | White beam mirror Y | Monochromator X | KB mirror sys HFMX-downstream | KB VFM fine pitch (readback only) | Alternative SSA dX (450mA) | HRM1 Y | Divergence-defining aperture Y(450mA) | In-air measurement Scan Y | Secondary Focusing working distance Z1 | Slits Guard SG1 dX (450mA) | Slits Guard SG2 dX (450mA)
Channel 3 | White beam mirror Pitch | Monochromator DCM.Offset2 | KB mirror sys HFMY-upstream | Beam viewing screen #3 Y | Alternative SSA Y (450mA) | HRM2 Y | Divergence-defining aperture dX (450mA) |  | WAXS detector1 X | Slits Guard SG1 Y (450mA) | Slits Guard SG2 Y (450mA)
Channel 4 | Multilayer monochromator theta1 | Monochromator DCM.Pitch2 | KB mirror sys HFMY-downstream | Beam viewing screen #4 Y | Alternative SSA dY (450mA) | Fine-resolution scanning (3000 mA) | Divergence-defining aperture dY (450mA) | microscope X | WAXS detector1 Y | Slits Guard SG1 dY (450mA) | Slits Guard SG2 dY (450mA)
Channel 5 | Multilayer monochromator X1 | Monochromator DCM.Roll2 | KB mirror sys VFMX | Slit S0 X-inboard | Attenuator X1 (75mA) | Transfocator X-upstream |  | microscope Y | WAXS detector1 Z | SAXS beam stop R1 (130mA) | SAXS beam stop X (170mA)
Channel 6 | Multilayer monochromator theta2 | Monochromator CCM.finePitch | KB mirror sys VFMY-upstream | Slit S0 X-outboard | Attenuator X2 (75mA) | Transfocator Y-upstream |  | SAXS detector X | WAXS detector2 X | SAXS beam stop R2 (130mA) | SAXS beam stop Y (170mA)
Channel 7 | Multilayer monochromator X2 | Monochromator CCM.finePitch (readback only) ?? | KB mirror sys VFMY-downstream | Slit S0 Y-top | Attenuator X3 (75mA) | Transfocator X-downstream | microscope polarizer/analyzer (130 mA) | SAXS detector Y | WAXS detector2 Y | SAXS beam stop R3 (130mA)
Channel 8 |  |  |  | Slit S0 Y-bottom | HRM2 theta | Transfocator Y-downstream |  | SAXS detector Z | WAXS detector2 Z | parfocus (250 mA)
Note on D12: VFM will now have only one horizontal translation

## Motor Channels
BEAMLINE: | Lix
| Totals → | 63 | 9 | 16 | >> This Subtotal does NOT include in the Front End. <<
Device | Motion Axis | Stepper Motor | Piezo | Other | Maximum | Minimum | Units | Notes | 5A amplifier | 0.1A-1A amplifier | 25mA-250mA
Front End
XBPM1 | X | 1 | [5 empty] | , ????
| Y | 1 | [5 empty] | , ????
XY Slits | X-inboard | 1 | [5 empty] | , ????
| X-outboard | 1 | [5 empty] | , ????
| Y-top | 1 | [5 empty] | , ????
| Y-bottom | 1 | [5 empty] | , ????
| Subtotal | 6 | 0 |  |  |  |  | >> This Subtotal NOT included in the Total above. <<
Hutch A (FOE)
Beam-viewing screen #1 | [7 empty] | Pneumatic
White beam mirror | X | 1 | [5 empty] | Inboard and Outboard, ~2A | 1
| Y | 1 | [5 empty] | Elevation, ~2A | 1
| Pitch | 1 | [5 empty] | ~2A | 1
| finePitch |  | 1
Beam-viewing screen #2 | Y | [6 empty] | Pneumatic
Multilayer monochromator | theta1 | 1 | [5 empty] | ~2A | 1
| X1 | 1 | [5 empty] | ~2A | 1
| theta2 | 1 | [5 empty] | ~2A | 1
| X2 | 1 | [5 empty] | ~2A | 1
| roll2 |  |  | 1 |  |  |  | likely something like PicoMotor
| finePitch2 |  | 1
Si(111) monochromator (DCM and CCM) | theta | 1 | [5 empty] | Bragg angle, ~2A | 1
| X | 1 | [5 empty] | switching between the DCM and the CCM, ~2A | 1
| DCM.Offset2 | 1 | [5 empty] | distance bwtween the two crystals, ~1A |  | 1
| DCM.Pitch2 | 1 | [5 empty] | second crystal pitch, ~1A |  | 1
| DCM.Roll2 | 1 | [5 empty] | second crystal roll, ~1A |  | 1
| DCM.finePitch2 |  | 1 |  |  |  |  | second crystal fine pitch, piezo
| CCM.finePitch | 1 | [5 empty] | ~1A |  | 1
Beam viewing screen #3 | Y | 1 | [5 empty] | ~2A | 1
Slits (S0) | X-inboard | 1 | [5 empty] | ~2A | 1
| X-outboard | 1 | [5 empty] | ~2A | 1
| Y-top | 1 | [5 empty] | ~2A | 1
| Y-bottom | 1 | [5 empty] | ~2A | 1
KB mirror system | HFMX-upstream | 1 | [5 empty] | 2.5A | 1
| HFMX-downstream | 1 | [5 empty] | 2.5A | 1
| HFMY-upstream | 1 | [5 empty] | 2.5A | 1
| HFMY-downstream | 1 | [5 empty] | 2.5A | 1
| VFMX | 1 | [5 empty] | 2.5A | 1
[8 empty] | VFMX-downstream has been removed | 1
| VFMY-upstream | 1 | [5 empty] | 2.5A
| VFMY-downstream | 1 | [5 empty] | 2.5A | 1
| fine pitch adjustments |  | 2 |  |  |  |  | PI piezo actuator, controller to be provided by us
Beam viewing screen #4 | Y | 1 | [5 empty] | ~2A | 1
| Subtotal | 26 | 5 | 1 |  |  |  | Total axis: | 26 |  | 0
Hutch B (SOE)
BPM | X |  |  | 1 |  |  |  | SmarAct motor
SSA | dX |  | 1 |  |  |  |  | horizontal opening, PiezoSystem Jena
| dY |  | 1 |  |  |  |  | vertical opening, PiezoSystem Jena
Attenuator | X1, X2, X3 | 3 | [5 empty] | low-current stepper, 75mA |  |  | 3
Video beam monitor | focus |  |  | 1 |  |  |  | open-loop PicoMotor,
| zoom |  |  | 1 |  |  |  | open-loop PicoMotor, not essential
Shutter/Scintilator | theta |  |  | 1 |  |  |  | Bistable solenoid, have driver, requires 1 TTL input (PLC?), also need to coordinate with detector trigger (software or TTL)
Alternative SSA | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
| Subtotal | 7 | 2 | 4 |  |  |  | Total axis | 0 | 4 | 3
Hutch C (Experimental Station)
Secondary Focusing working distance | Z | 1 | [5 empty] | Adjust working distance of the transfocator in vacuum, ~2A | 1
| Z1 | 1 | [5 empty] | Adjust the location of the SF chamber, ~2A | 1
Transfocator | X-upstram | 1 | [5 empty] | Kohzu AXA04A-R101-1F Motor PKP214D06B 2ph .6A | 1
| Y-upstream | 1 | [5 empty] | Kohzu AZA05A-W1-1C Motor: PKP214D06B 2ph .6A | 1
| X-downstream | 1 | [5 empty] | ~1.5A | 1
| Y-downstream | 1 | [5 empty] | ~1.5A | 1
| CRLgroup1-9 |  |  | 9 |  |  |  | brushless DC motors, each motor require two inputs (enable and direction) from the PLC
Divergence-defining aperture (DDA) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
HRM1 | theta | 1 | [5 empty] | Kohzu SA04B-RT-2N Motor: PKP223D15B 2ph, 1.5A | 1
| Y | 1 | [5 empty] | Kohzu; ZA07A-X1-1T Motor: PKP243D15B,2ph 1.5A | 1
HRM2 | theta, tilt | 1 | [5 empty] | Tilt: Newport URS50BPPV6 15 deg/step, 0.16A |  |  | 1
| Y | 1 | [5 empty] | Kohzu; ZA07A-X1-1T Motor: PKP243D15B,2ph 1.5A | 1
Guard slits 1 (Sg1) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
BIM1 | Y |  |  | 1 |  |  |  | SmarAct, moves a diamond quad-BPM
Guard slits 2 (Sg2 for grazing incidence) | X, Y, dX, dY | 4 | [5 empty] | low-current stepper, 450mA |  | 4
Guard slits 2 (Sg2 for transmission) | dX, dY |  |  | 2 |  |  |  | SmarAct
| Subtotal | 22 | 0 | 12
Microscope | zoom, focus | 2 | [5 empty] | , ????
| polarizer/analyzer | 1 | [5 empty] | , ????
| alignment X and Y | 2 | [5 empty] | , ????
| position X , Y, focus | 3 | [5 empty] | , ????
In-air measurement | scanX, scanY | 2 | [5 empty] | scanX, scanY, stepper 670 mA/phase |  | 2
| finescanX, finescanY |  | 2 |  |  |  |  | pizeo scanning stage from PI
| azimuthal | 1 | [5 empty] | stepper 2 phase, 1.8 deg/step NEMA23 ~1A | 1
in-vacuum GISAXS/GID with hexapod | X, Y, Z, U, V, W, phi |  |  | 7 |  |  |  | already have controller from miCos/PI
GISAXS/GID | X,Y | 2 | [5 empty] | stepper, probably size 23, ~1A | 2
| tiltX, tiltY |  |  | 2 |  |  |  | Newport
WAXS detectors | X1, Y1, X1, X2, Y2, Z2 | 6 | [5 empty] | ~2.5A | 6
WAXS beamstop | X, Y, Z | 3 | [5 empty] | SmarAct
SAXS detector | X, Y, Z | 3 | [5 empty] | large steppers | 3
SAXS beam stop | R1, R2, R3 | 3 | [5 empty] | low-current steppers, ~130mA |  |  | 3
| X, Y | 2 | [5 empty] | small servo motors, 170mA |  |  | 2
| tilt1, tilt2 |  |  | 2 |  |  |  | Open-loop PicoMotors
| Subtotal | 30 | 2 | 11 |  |  |  | Total minus ????: | 21 | 14 | 6
| total | 85 | 9 | 28 | 122

## RackLayout
These are rackes on the roofs
| RG A-1 |  |  | RG-A1 back |  |  | RG-A2 front |  |  |  | RG-B1 |  |  |  |  | RG-C1 |  |  | RG-C2
42 |  | 42 | 42 |  | 42 | 42 |  |  | 42 | 42 |  |  | 42 |  | 42 |  | 42 | 42 |  | 42
41 |  | 41 | 41 | fiber runs to this switch | 41 | 41 |  |  | 41 | 41 |  |  | 41 |  | 41 |  | 41 | 41 |  | 41
40 |  | 40 | 40 | switch | 40 | 40 | TEMP. CONTROL |  | 40 | 40 | TEMP. CONTROL |  | 40 |  | 40 | TEMP. CONTROL | 40 | 40 | TEMP. CONTROL | 40
39 |  | 39 | 39 | cable management | 39 | 39 | SPACER |  | 39 | 39 | MOXA |  | 39 |  | 39 | fiber patch pannel | 39 | 39 | fiber patch pannel | 39
38 | IOC | 38 | 38 |  | 38 | 38 | MOXA |  | 38 | 38 | SPACER |  | 38 |  | 38 | switch x48ports | 38 | 38 | MOXA | 38
37 | MOXA | 37 | 37 |  | 37 | 37 | SPACER |  | 37 | 37 | UPS |  | 37 |  | 37 | cable management | 37 | 37 |  | 37
36 | SPACER | 36 | 36 |  | 36 | 36 | #1 GAUGE CONTROL | #2 GAUGE CONTROL | 36 | 36 |  |  | 36 |  | 36 | cable management | 36 | 36 |  | 36
35 | UPS | 35 | 35 |  | 35 | 35 |  |  | 35 | 35 | SPACER |  | 35 |  | 35 | reserve switch future | 35 | 35 |  | 35
34 |  | 34 | 34 |  | 34 | 34 | SPACER |  | 34 | 34 |  |  | 34 |  | 34 | cable managent future | 34 | 34 |  | 34
33 | SPACER | 33 | 33 |  | 33 | 33 | #3 GAUGE CONTROL | #4 GAUGE CONTROL | 33 | 33 |  |  | 33 |  | 33 | cable managent future | 33 | 33 | Archiver IOC | 33
32 | HV-ADAPTOS 19' wide | 32 | 32 |  | 32 | 32 |  |  | 32 | 32 |  |  | 32 |  | 32 | IOC | 32 | 32 |  | 32 | ioc should be 32 or higher because of the ppwer strips
31 |  | 31 | 31 |  | 31 | 31 | SPACER |  | 31 | 31 |  |  | 31 |  | 31 | from B/C hutch patch pannel x12 | 31 | 31 | BNC/lemo patch panel x16 | 31
30 |  | 30 | 30 |  | 30 | 30 | VACUUM PLC |  | 30 | 30 |  |  | 30 |  | 30 | network patch pannel x24 | 30 | 30 | TIMING SYSTEM VME | 30 |  | Inside B hutch
29 |  | 29 | 29 |  | 29 | 29 |  |  | 29 | 29 |  |  | 29 |  | 29 | SPACER | 29 | 29 |  | 29 |  | BNC/lemo patch panel x8
28 | TEMP. CONTROL | 28 | 28 |  | 28 | 28 |  |  | 28 | 28 |  |  | 28 |  | 28 | UPS | 28 | 28 |  | 28
27 | beam stabilization | 27 | 27 |  | 27 | 27 |  |  | 27 | 27 |  |  | 27 |  | 27 |  | 27 | 27 |  | 27
26 |  | 26 | 26 |  | 26 | 26 | SPACER |  | 26 | 26 | #1 GAUGE CONTROL | #2 GAUGE CONTROL | 26 | ????? | 26 | SPACER | 26 | 26 |  | 26
25 |  | 25 | 25 |  | 25 | 25 | ION PUMP CONTROL #1 |  | 25 | 25 |  |  | 25 | ?????? | 25 | PPU | 25 | 25 |  | 25
24 |  | 24 | 24 |  | 24 | 24 |  |  | 24 | 24 | SPACER |  | 24 |  | 24 | P1M control | 24 | 24 | VME shelf | 24
23 |  | 23 | 23 |  | 23 | 23 |  |  | 23 | 23 | EPS BEAMLINE PLC |  | 23 |  | 23 | SPACER | 23 | 23 | SPACER | 23 |  | Inside C hutch
22 | SPACER | 22 | 22 |  | 22 | 22 | SPACER |  | 22 | 22 |  |  | 22 |  | 22 | P300K control | 22 | 22 | DELTA-TAU PSU | 22 |  | BNC/lemo patch panel x8
21 | DELTA-TAU PSU | 21 | 21 |  | 21 | 21 | ION PUMP CONTROL #2 |  | 21 | 21 |  |  | 21 |  | 21 | P300K control | 21 | 21 | SPACER | 21
20 | SPACER | 20 | 20 |  | 20 | 20 |  |  | 20 | 20 |  |  | 20 |  | 20 |  | 20 | 20 | DELTA-TAU#11 low current | 20
19 | DELTA-TAU#4 | 19 | 19 |  | 19 | 19 |  |  | 19 | 19 | SPACER |  | 19 |  | 19 |  | 19 | 19 |  | 19
18 |  | 18 | 18 |  | 18 | 18 | SPACER |  | 18 | 18 | DELTA-TAU PSU |  | 18 |  | 18 |  | 18 | 18 |  | 18
17 |  | 17 | 17 |  | 17 | 17 | ION PUMP CONTROL #3 |  | 17 | 17 | SPACER |  | 17 |  | 17 |  | 17 | 17 | SPACER | 17
16 | SPACER | 16 | 16 |  | 16 | 16 |  |  | 16 | 16 | DELTA-TAU#7 low current |  | 16 |  | 16 |  | 16 | 16 | DELTA-TAU#10 low current | 16
15 | DELTA-TAU#3 | 15 | 15 |  | 15 | 15 |  |  | 15 | 15 |  |  | 15 |  | 15 |  | 15 | 15 |  | 15
14 |  | 14 | 14 |  | 14 | 14 | SPACER |  | 14 | 14 |  |  | 14 |  | 14 |  | 14 | 14 |  | 14
13 |  | 13 | 13 |  | 13 | 13 | ION PUMP CONTROL #4 |  | 13 | 13 | SPACER |  | 13 |  | 13 |  | 13 | 13 | SPACER | 13
12 | SPACER | 12 | 12 |  | 12 | 12 |  |  | 12 | 12 | DELTA-TAU#6 |  | 12 |  | 12 | Chillers for Pilatus detectors | 12 | 12 | DELTA-TAU#9 | 12
11 | DELTA-TAU#2 | 11 | 11 |  | 11 | 11 |  |  | 11 | 11 |  |  | 11 |  | 11 |  | 11 | 11 |  | 11
10 |  | 10 | 10 |  | 10 | 10 | SPACER |  | 10 | 10 |  |  | 10 |  | 10 |  | 10 | 10 |  | 10
9 |  | 9 | 9 |  | 9 | 9 | ION PUMP CONTROL #5 |  | 9 | 9 | SPACER |  | 9 |  | 9 |  | 9 | 9 | SPACER | 9
8 | SPACER | 8 | 8 |  | 8 | 8 |  |  | 8 | 8 | DELTA-TAU#5 low current |  | 8 |  | 8 |  | 8 | 8 | DELTA-TAU#8 | 8
7 | DELTA-TAU#1 | 7 | 7 |  | 7 | 7 |  |  | 7 | 7 |  |  | 7 |  | 7 |  | 7 | 7 |  | 7
6 |  | 6 | 6 |  | 6 | 6 | SPACER |  | 6 | 6 |  |  | 6 |  | 6 |  | 6 | 6 |  | 6 |  | Inside B hutch
5 |  | 5 | 5 |  | 5 | 5 | RGA controller |  | 5 | 5 | network patch pannel x12 |  | 5 |  | 5 |  | 5 | 5 | network patch pannel x12 | 5 |  | network patch pannel x4
4 |  | 4 | 4 |  | 4 | 4 |  |  | 4 | 4 |  |  | 4 |  | 4 | water tray for leaks | 4 | 4 |  | 4
3 |  | 3 | 3 |  | 3 | 3 |  |  | 3 | 3 |  |  | 3 |  | 3 |  | 3 | 3 |  | 3 |  | Inside C hutch
2 |  | 2 | 2 |  | 2 | 2 |  |  | 2 | 2 |  |  | 2 |  | 2 |  | 2 | 2 |  | 2 |  | network patch pannel x8
1 |  | 1 | 1 |  | 1 | 1 |  |  | 1 | 1 |  |  | 1 |  | 1 |  | 1 | 1 |  | 1
controller for WAXS positioning stages
UPS ~4kW
Note on Q7: SmarAct controller ??, pico motor controller?? BPM electronics??

## RackLayout_Floor
Thise are racks on the floor.
| RG Floor-1 |  |  | RG-Floor2
42 |  | 42 | 42 |  | 42
41 |  | 41 | 41 |  | 41
40 | TEMP. CONTROL | 40 | 40 |  | 40
39 | fiber patch pannel | 39 | 39 |  | 39
38 | SPACER | 38 | 38 |  | 38
37 | reserve future | 37 | 37 |  | 37
36 | reserve future | 36 | 36 |  | 36
35 | reserve future | 35 | 35 |  | 35
34 | reserve future | 34 | 34 |  | 34
33 | path panel | 33 | 33 |  | 33
32 | cable management | 32 | 32 |  | 32
31 | SPACER | 31 | 31 |  | 31
30 | GPU Server #1 | 30 | 30 |  | 30
29 | SPACER | 29 | 29 |  | 29
28 | GPU Server #2 | 28 | 28 |  | 28
27 | SPACER | 27 | 27 |  | 27
26 | STORAGE SERVER #1 | 26 | 26 |  | 26
25 | SPACER | 25 | 25 |  | 25
24 | STORAGE SERVER #2 | 24 | 24 |  | 24
23 |  | 23 | 23 |  | 23
22 |  | 22 | 22 |  | 22
21 |  | 21 | 21 |  | 21
20 |  | 20 | 20 |  | 20
19 | SPACER | 19 | 19 |  | 19
18 | Storage #1 | 18 | 18 |  | 18
17 |  | 17 | 17 |  | 17
16 | SPACER | 16 | 16 |  | 16
15 | Storage #2 | 15 | 15 |  | 15
14 |  | 14 | 14 |  | 14
13 |  | 13 | 13 |  | 13
12 |  | 12 | 12 |  | 12
11 |  | 11 | 11 |  | 11
10 |  | 10 | 10 |  | 10
9 |  | 9 | 9 |  | 9
8 |  | 8 | 8 |  | 8
7 |  | 7 | 7 |  | 7
6 |  | 6 | 6 |  | 6
5 |  | 5 | 5 |  | 5
4 |  | 4 | 4 |  | 4
3 |  | 3 | 3 |  | 3
2 | Power Supply | 2 | 2 | Power Supply | 2
1 |  | 1 | 1 |  | 1

## A Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:4}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:4 |  | NA | NA | 65 | 64
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| 94 | 93 | 92 | 91 | 108 | 95 | NA | NA
CONTROLLER |  | XF:16IDA-CT{MC:3}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:3 |  | 60 | 59 | 58 | 57
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 63 | 62 | 61
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | 90 | 89 | 88 | 87 | 86 | 85 | 84
CONTROLLER |  | XF:16IDA-CT{MC:2}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:2 |  | 54 | 53 | 52 | 51
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 56 | NA | 55
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | NA | 83 | 82 | 81 | 80 | 79 | 78
CONTROLLER |  | XF:16IDA-CT{MC:1}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:1 |  | 118 | 68 | 67 | 66
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | 121 | 120 | 119
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | 117 | 116 | 115 | 114 | 98 | 97 | 96 | This cables had to be separately pulled

## B Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:7}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:7 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | free | free
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | C-106 | free | free | C-18 | C-17 | C-16 | C-15
CONTROLLER |  | XF:16IDA-CT{MC:6}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:6 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-14 | C-13 | C-12 | C-11 | C-24 | C-21 | C-20 | C-19
CONTROLLER |  | XF:16IDA-CT{MC:5}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:5 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | C-23 | B-8 | B-7 | B-6
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-22 | B-10 | B-9 | B-5 | B-23 | B-22 | B-12 | B-11

## C Hutch MC terminations
CONTROLLER |  | XF:16IDA-CT{MC:11}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:11 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | free | free | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| free | free | C-52 | C-51 | C-30 | C-29 | C-28 | C-27
CONTROLLER |  | XF:16IDA-CT{MC:10}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:10 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-33 | C-67 | C-66 | C-65 | C-26 | C-25 | C-10 | C-9
CONTROLLER |  | XF:16IDA-CT{MC:9}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
C-104 C-105
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:9 |  | NA | NA | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | NA | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-41 | C-40 | C-39 | C-38 | C-37 | C-36 | C-8 | C-7
CONTROLLER |  | XF:16IDA-CT{MC:8}
5 V EXT ENC | 24 V LOGIC/IO | 48 V MOTOR | WATCHDOG/ABORT |  |  |  |  | GPIO
C-107
|  |  | ENCODER 4 | ENCODER 3 | ENCODER 2 | ENCODER 1
| MC:8 |  | NA | free | NA | NA
|  |  | ENCODER 8 | ENCODER 7 | ENCODER 6 | ENCODER 5
|  |  | C-45 C-46 | NA | NA | NA
| AMP 8 | AMP 7 | AMP 6 | AMP 5 | AMP 4 | AMP 3 | AMP 2 | AMP 1
| C-50 | C-49 | C-48 | C-32 | C-31 | free | C-35 | C-34
Special wiring